# Friday 11th Sep 2020	(VD & AD)

# Set up after zombie apocalypse

Refamiliarising with the chamber

11.10	Rotary pump switched on

11:13	Turbo pump switched on

Pressure came down quick – system clean ;) (10E-6 mbar within 1h = no problem)

11:43	Chamber pressure 4 x10-6 mbar

14:51 	p=5x10-7 mbar

15:30	MCT detector cooled

15:45	FTIR focusing mirror & detector alignment done Intensity 4.71

# Tuesday 15th Sep 2020	(VD & AD)

08:50	p=2x10-8 mbar; T=room temp

Clean substrate: turned the heater on (16V) and watched pressure rise until it reached a turning point at 3x10-6 mbar; held for 2-3 min and turned the heater off.

9:05 	p = 5*10-8 mbar

9:15 	p = 4 *10-8 mbar

9:17 	Cryo compressor turned on

10:18 	T = 103K / P = 5 * 10-9 mbar

10:20 	Start cooling MCT detector

10:50 	New rotary pump installed – connected to gas line

10:58 	Pump (GL) turned on

11:00 	T = 22.3 K (base temperature ?) P = 3.50 10-9 mbar

11:05 	Deionised Water connected and F-P-T carried out x 3

13:15 	Laser aligned signal 321 mV

14:22 	P

Background settings:

Sample setting

## Background scan(s) #1

15:00	BG20200915_01

Location : C:\Documents and Settings\MolPhys\My Documents\Vincent\2020_09_15
512 scans	1 cm-1 res	signal 
p = 6 x 10-10 mbar	T = 20.5 K   Wrong values (check above for partial info)

## Deposition #1: H2O @ 22 K

15:30 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 22.1K

- Initial pressure: 3*10-9 mbar

- Gas cell pressure: 9.5 torr

- Deposition pressure: 1 – 1.2 10E-7 mbar

- Laser signal: 325 mV

- Deposition time: 20 min

Something funny happened with the laser signal output – not much change on mV scale:

16:13	ASW_2020_09_15_0001

512 scans	1 cm-1 res	signal 4.86
p = 6.7 x 10-10 mbar	T = 20.5 K

PID values changed p = 2.5 I = 5.0 d = 0.0

## Anneal to 130 K

17:06 Set Temperature 130K

Pressure chamber increased slightly (to monitor ?)

17:32	ASW_2020_09_15_0002

512 scans	1 cm-1 res	signal 4.84
p = ?? mbar	T = 130 K

## Anneal to 140 K

17:50 	Set Temperature 140K

18:07	ASW_2020_09_15_0003

512 scans	1 cm-1 res	signal 4.84
p = ??mbar	T = 140 K

## Anneal to 160 K

18:25 Set Temperature 160K

18:48 	ASW_2020_09_15_0004  512 scan, resolution 1, signal: 4.90

512 scans	1 cm-1 res	signal 4.90
p = ?? mbar	T = 160 K

19:12 Heater set up manual 0.0 V and cryostat turned off.

Why is 160K such a drop in Absorbance?

Great day,

# Wednesday 16th Sep 2020	(VD & AD)

09:00	p= 5x10-8 mbar

Heater on Auto to 300 K

09:30 	MCT detector cooled; cryo on

Plan for the day:

- Deposit same as yesterday – 1-1.2 x 10-7 mbar for 20 min

- Scan at 20, 40, 60, 80 (see dangling bonds disappear), 100, 120, 130, 140, 160…

- If there is time, heat to 200 to remove water, cool to 130-160 and deposit

10:45	p=4x10-9 mbar		T=70 K and dropping

## Background scan(s) #1

11:04	BG20200916_01

512 scans	res = 1 cm-1	signal = 5.02
p = 3.5 10-9 mbar	T = 22.2 K

Red scan = 16/09     Green = 15/09  Less water, sample cleaner

## Deposition #2: H2O @ 22 K

12:00 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 22.1 K

- Initial pressure: 3*10-9 mbar

- Initial gas cell pressure: 7.64 Torr

- Deposition pressure:

- Laser signal: 324.9 mV

- Deposition time: 20 min

- pressure after deposition 6.7*10-9 mbar

- final gas cell pressure = 6.4 Torr

voltage/mV

320.0

321.0

322.0

323.0

324.0

325.0

326.0

0

200

400

600

800

1000

1200

voltage/mV

Pressure spike!

12:27	laser Off

12:30	ASW_2020_09_16_0001

512 scans	1 cm-1 res	signal 4.94
p = 6.9 x 10-9 mbar	T = 22.1 K

13:30	ASW_2020_09_16_0002 – repeat scan after 1 hour

512 scans	1 cm-1 res	signal 4.92
p = 3.1 x 10-9 mbar	T = 22.1 K

## Annealing to 40 K

All files: ASW_2020_09_16_000# (see tables) 512 scans; 1 cm-1 res

Temperature recorder as a function of time – EXCEL file: Temperature-steps-2020_09_16.xlsx

| time | File | Start Temp/K | End Temp |  | p/mbar | signal |
| --- | --- | --- | --- | --- | --- | --- |
| 14 :00 | _0003 | 20 | 39.9 | Warm-up | 2x10-8 – 3.3x10-9 | 4.87 |
| 14 :15 | _0004 | 39.9 | 40.0 | Warm-up | 3.3x10-9 | 4.87 |
| 14 :30 | _0005 | 40.0 | 40.0 | isothermal | 3.0x10-9 | 4.87 |

-0.00

0.02

0.04

0.06

0.08

0.10

0.12

0.14

0.16

0.18

0.20

0.22

0.24

0.26

0.28

0.30

1000

1500

2000

2500

3000

3500

4000

Wavenumbers (cm-1)

20-40 K warm-up

22.1 K isothermal

22.1 K deposition

## Annealing to 60 K

| 14 :45 | _0006 | 40.0 | 59.8 | Warm-up | 4.8x10-8 – 4x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 15 :00 | _0007 | 59.8 | 60 | Warm-up | 4.0x10-8 – 3.3x10-9 | 4.84 |
| 15 :15 | _0008 | 60 | 60 | isothermal | 3.3x10-8 – x10-9 | 4.84 |

## Annealing to 80 K

| 15 :30 | _0009 | 60 | ## | Warm-up | 4.8x10-8 – 4x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 15 :45 | _0010 | ## | 80 | Warm-up | 4.0x10-8 – 3.3x10-9 | 4.84 |
| 16 :00 | _0011 | 80 | 80 | isothermal | 3.3x10-8 – x10-9 | 4.84 |

## Annealing to 100 K

| 16:15 | _0012 | 80 | ## | Warm-up | 4.8x10-8 – 4x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :30 | _0013 | ## | 100 | Warm-up | 4.0x10-8 – 3.3x10-9 | 4.84 |
| 16 :45 | _0014 | 100 | 100 | isothermal | 3.3x10-8 – x10-9 | 4.84 |

## Annealing to 120 K

| 17 :00 | _0015 | 100 | ## | Warm-up | 4.8x10-8 – 4x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :15 | _0016 | ## | 120 | Warm-up | 4.0x10-8 – 3.3x10-9 | 4.84 |
| 17 :30 | _0017 | 120 | 120 | isothermal | 3.3x10-8 – x10-9 | 4.84 |

## Annealing to 130 K

| 17 :45 | _0018 | 120 | ## | Warm-up | 4.8x10-8 – 4x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :00 | _0019 | ## | 130 | Warm-up | 4.0x10-8 – 3.3x10-9 | 4.84 |
| 18 :15 | _0020 | 130 | 130 | isothermal | 3.3x10-8 – x10-9 | 4.84 |

## Annealing to 140 K

| 18:30 | _0021 | 130 | ## | Warm-up | 4.8x10-8 – 4x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :45 | _0022 | ## | 140 | Warm-up | 4.0x10-8 – 3.3x10-9 | 4.84 |
| 19 :00 | _0023 | 140 | 140 | isothermal | 3.3x10-8 – x10-9 | 4.84 |

## Annealing to 160 K

| 19 :15 | _0024 | 140 | ## | Warm-up | 4.8x10-8 – 4x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 19 :30 | _0025 | ## | 160 | Warm-up | 4.0x10-8 – 3.3x10-9 | 4.84 |
| 19 :45 | _0026 | 160 | 160 | isothermal | 3.3x10-8 – x10-9 | 4.84 |

## Annealing to 180 K

| 20 :00 | _0027 | 160 | ## | Warm-up | 4.8x10-8 – 4x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 20 :15 | _0028 | ## | 180 | Warm-up | 4.0x10-8 – 3.3x10-9 | 4.84 |
| 20 :30 | _0029 | 180 | 180 | isothermal | 3.3x10-8 – x10-9 | 4.84 |

## Annealing to 200 K

| 20 :45 | _0030 | 180 | ## | Warm-up | 4.8x10-8 – 4x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 21 :00 | _0031 | ## | 200 | Warm-up | 4.0x10-8 – 3.3x10-9 | 4.84 |

To complete  P wrongs

# Thursday 17th Sep 2020	(VD & AD)

08:48	p= 2.25x10-8 mbar

Heater on Man at 0V  T = 275.4

08:50 	Cryo on

9:07	Laser turned on

9:26	MCT Detector cooled (ice was growing on the edges of the opening and I put the tape on top of it, is it an issue?)

10:12	T = 22K  Did wrong manip (put heater on auto)  T increase to 44K (quick flash heat)  put back heater on Man V = 0.0

10:42	Base P = 2.4 10E-9 	Base T = 21.7

## Background scan(s) #1

11:00	BG20200917_01

512 scans	res = 1 cm-1	signal = 5.15
p = 2.3 10-9 mbar	T = 21.7 K

11 :20	Refill gas cell procedure (0.642 Torr initial pressure)

Gas cell valve open

Pump valve close

{...}

Final Pressure : 0.759 Torr

11:33	Head rotated Laser signal = 325.4 mV

## Deposition #3: H2O @ 22 K

11:40 	2 min @ 1x10-6 mbar H2O

- Initial Temperature: 21.7  K

- Initial pressure: 2.2 10-9 mbar

- Initial gas cell pressure: 7.56 Torr

- Deposition pressure: 1.04 10-6 – 9.8 10-7

- Laser signal:  325.4 mV

- Deposition time: 2 min

- pressure after deposition 8.9*10-9 mbar

- final gas cell pressure = 6.32 Torr

11:47	laser Off

11:48	Head rotated

11:52	ASW_2020_09_17_0001

512 scans	1 cm-1 res	signal 5.27
p = 5.7 x 10-9 mbar	T = 21.7 K

12:08	ASW_2020_09_17_0002

512 scans	1 cm-1 res	signal 5.26
p = 3.5 x 10-9 mbar	T = 21.7 K

13 :17	P = 2.3 10-9 mbar	T = 21.6 K

## Annealing to 40 K

All files: ASW_2020_09_17_000# (see tables) 512 scans; 1 cm-1 res

Temperature recorder as a function of time – EXCEL file: Temperature-steps-2020_09_17.xlsx

| time | File | Start Temp/K | End Temp |  | p/mbar | signal |
| --- | --- | --- | --- | --- | --- | --- |
| 14 :00 | _0003 | 21.6 | 39.9 | Warm-up | 3x10-8 – 2.5x10-9 |  |
| 14 :15 | _0004 | 39.9 | 40.0 | Warm-up | 2.5x10-9 |  |
| 14 :30 | _0005 | 40.0 | 40.0 | isothermal | 2.3x10-9 | 5.24 |

## Annealing to 60 K

| 14 :45 | _0006 | 40.0 | 59.8 | Warm-up | 8x10-8 – 3.6x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 15 :00 | _0007 | 59.8 | 60 | Warm-up | 3.1x10-9 | 5.23 |
| 15 :15 | _0008 | 60 | 60 | isothermal | 2.7 x10-9 | 5.22 |

## Annealing to 80 K

| 15 :30 | _0009 | 60 | 79.7 | Warm-up | 2.5x10-8 – 3.1x10-9 | 5.21 |
| --- | --- | --- | --- | --- | --- | --- |
| 15 :45 | _0010 | 79.8 | 79.9 | Warm-up | 2.9 x10-9 | 5.21 |
| 16 :00 | _0011 | 79.9 | 80 | isothermal | 2.8 x10-9 | 5.21 |

## Annealing to 100 K

| 16:15 | _0012 | 80 | 99.8 | Warm-up | 5.7x10-8 – 4.1x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :30 | _0013 | 99.9 | 100 | Warm-up | 3.6 x10-9 | 5.20 |
| 16 :45 | _0014 | 100 | 100 | isothermal | 3.0 x10-9 | 5.20 |

## Annealing to 120 K

| 17 :00 | _0015 | 100 | 119.7 | Warm-up | 7.7x10-8 – 4.8x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :15 | _0016 | 119.7 | 120 | Warm-up | 3.6 x10-9 | 5.19 |
| 17 :30 | _0017 | 120 | 120 | isothermal | 3.1 x10-9 | 5.19 |

## Annealing to 130 K

| 17 :45 | _0018 | 120 | 129.8 | Warm-up | 2.6x10-8 – 7.7x10-9 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :00 | _0019 | 129.8 | 130 | Warm-up | 7.1x10-8 | 5.18 |
| 18 :15 | _0020 | 130 | 130 | isothermal | 7.0x10-8 | 5.17 |

## Annealing to 140 K

| 18:30 | _0021 | 130 | 139.8 | Warm-up | 8x10-8 – 5.9x10-8 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :45 | _0022 | 139.8 | 140 | Warm-up | 5.3x10-8 | 5.17 |
| 19 :00 | _0023 | 140 | 140 | isothermal | 4.6x10-8 | 5.17 |

## Annealing to 160 K

| 19 :15 | _0024 | 140 | 159.7 | Warm-up | 1.1x10-7 – 3.4x10-6 | 5.15 |
| --- | --- | --- | --- | --- | --- | --- |
| 19 :30 | _0025 | 159.8 | 160 | Warm-up | 2.9x10-6 | 5.16 |
| 19 :45 | _0026 | 160 | 160 | isothermal | 1.2x10-6 | 5.04 |

## Annealing to 180 K

| 20 :00 | _0027 | 160 | ## | Warm-up | ## | 5.02 |
| --- | --- | --- | --- | --- | --- | --- |
| 20 :15 | _0028 | ## | 180 | Warm-up |  | 5.01 |

20:30	Gas cell valve open for pumping

20:30	Heater 0.0 V

20:30	Cryo off

# Monday 21th Sep 2020	(VD)

08:42	p= x10-8 mbar; T=room temp (278.5)

Clean substrate: turned the heater on (16V) and watched pressure rise until it reached a turning point at ~ 10-7 mbar; held for 2-3 min and turned the heater off.

8:51 	p = 2.1*10-8 mbar

8:55	p = 1.7 *10-8 mbar

8:55	Cryo-on, Laser on

9:34	MCT cooled

## Background scan(s) #1

10:56	BG20200921_01

512 scans	res = 1 cm-1	signal = 5.11
               p = 1.3 * 10-9 mbar          T = 21.8K

11 :20	Refill gas cell procedure (GAS CELL LEFT PUMPING OVER THE WEEK-END)

Gas cell valve open

Pump valve close

{...}

Final Pressure : 0.73 Torr

11:30 Head rotated laser signal = 325.3

## Deposition #4: H2O @ 22 K

12:00 	3h20 @ 1x10-8 mbar H2O

- Initial Temperature: 21.8  K

- Initial pressure: 1.3 10-9 mbar

- Initial gas cell pressure: 6.92 Torr

- Initial gas cell pressure: 5.87 Torr

- Deposition pressure: 1 * 10-8 mbar

- Laser signal:  325.3 mV

- Deposition time: 3h20 min

15:20	Needle valve closed – laser turned off

15:22	P = 2.1 * 10-9 mbar T = 21.8

15:25	ASW_2020_09_21_0001

512 scans	1 cm-1 res	signal 5.10
p = 1.9 x 10-9 mbar	T = 21.8  K

15:57	ASW_2020_09_21_0002

512 scans	1 cm-1 res	signal 5.08
p = 1.3 x 10-9 mbar	T = 21.8 K

## Annealing to 40 K

All files: ASW_2020_09_21_000# (see tables) 512 scans; 1 cm-1 res

Temperature recorder as a function of time – EXCEL file: Temperature-steps-2020_09_21.xlsx

| time | File | Start Temp/K | End Temp |  | p/mbar | signal |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :30 | _0003 | 21.7 | 39.9 | Warm-up |  |  |
| 16 :45 | _0004 | 39.9 | 40.1 | Warm-up |  | 5.06 |
| 17 :00 | _0005 | 40.1 | 40.0 | isothermal |  | 5.06 |

## Annealing to 60 K

| 17:15 | _0006 | 40.0 |  | Warm-up |  | 5.05 |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :30 | _0007 |  | 60 | Warm-up |  | 5.06 |
| 17 :45 | _0008 | 60 | 60 | isothermal |  | 5.05 |

## Annealing to 80 K

| 18 :00 | _0009 | 60 | ## | Warm-up |  | 5.04 |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :15 | _0010 | ## | 80 | Warm-up |  | 5.04 |
| 18 :30 | _0011 | 80 | 80 | isothermal |  | 5.03 |

## Annealing to 100 K

| 18:45 | _0012 | 80 | 99.8 | Warm-up |  | 5.03 |
| --- | --- | --- | --- | --- | --- | --- |
| 19:00 | _0013 | 99.9 | 100 | Warm-up |  | 5.03 |
| 19 :15 | _0014 | 100 | 100 | isothermal |  | 5.01 |

## Annealing to 120 K

| 19 :30 | _0015 | 100 | 119.7 | Warm-up |  | 5.01 |
| --- | --- | --- | --- | --- | --- | --- |
| 19 :45 | _0016 | 119.7 | 120 | Warm-up |  |  |
| 20 :00 | _0017 | 120 | 120 | isothermal |  | 5.01 |

## Annealing to 130 K

| 20 :15 | _0018 | 120 |  | Warm-up |  | 5.01 |
| --- | --- | --- | --- | --- | --- | --- |
| 20 :30 | _0019 | 129.8 | 130 | Warm-up |  | 5.01 |
| 20 :45 | _0020 | 130 | 130 | isothermal |  | 5.01 |

## Annealing to 140 K

| 21:00 | _0021 | 130 | 139.8 | Warm-up |  | 5.01 |
| --- | --- | --- | --- | --- | --- | --- |
| 21 :15 | _0022 | 139.8 | 140 | Warm-up |  | 5.02 |
| 21 :30 | _0023 | 140 | 140 | isothermal |  | 5.01 |

## Annealing to 160 K

| 21 :45 | _0024 | 140 | 159.7 | Warm-up |  | 5.01 |
| --- | --- | --- | --- | --- | --- | --- |
| 22 :00 | _0025 | 159.8 | 160 | Warm-up |  | 4.96 |
| 22 :15 | _0026 | 160 | 160 | isothermal |  | 4.97 |

## Annealing to 180 K

| 22 :30 | _0027 | 160 | ## | Warm-up | ## | 4.98 |
| --- | --- | --- | --- | --- | --- | --- |
| 22 :45 | _0028 | ## | 180 | Warm-up |  | 4.96 |

23 :00	Heater turned off

23 :05	Cryo turned of

# Thursday 24th Sep 2020	(VD)

9:25	T = 278.5	P = 7.80 * 10-9

9:30	Cryo turned on

10:21	T = 136.7	P = 1.9 * 10-9

11:43	T = 21.7 K	P = 9.8 10-10

12:04 T set up 100K

12:12 MCT detector cooled / Laser turned on

12:50	T= 100K		P= 1.9 * 10 -9

## Background scan(s) #1

13:40	BG20200924_01

512 scans	res = 1 cm-1	signal = 5.23
p = 1.5 10-9 mbar	T = 100.0 K

11 :20	Refill gas cell procedure (0.578 Torr initial pressure)

Pump valve close

{...}

Final Pressure : 0.753 Torr

11:33	Head rotated Laser signal = 325.2 mV

## Deposition #5: H2O @ 100K

11:40 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 100K

- Initial pressure: 1.5 10-9 mbar

- Initial gas cell pressure: 7.53 Torr

- Deposition pressure: 1.05 10-7 – 1.1 10-7

- Laser signal:  325 mV

- Deposition time: 20 min

- pressure after deposition 6.7*10-9 mbar

- final gas cell pressure = 6.64 Torr

14:32	laser Off

14:32	Head rotated

14:49	ASW_2020_09_24_0001

512 scans	1 cm-1 res	signal 5.34
p = 2.7 x 10-9 mbar	T = 100.0 K

15:05	ASW_2020_09_24_0002

512 scans	1 cm-1 res	signal 5.26
p =	T = 100 K

## Annealing to 120 K

| 16 :00 | _0003 | 100 | 119.7 | Warm-up |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :15 | _0004 | 119.7 | 120 | Warm-up |  | 5.30 |
| 16 :30 | _0005 | 120 | 120 | isothermal |  | 5.30 |

## Annealing to 130 K

| 16 :45 | _0006 | 120 |  | Warm-up |  | 5.30 |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :00 | _0007 |  | 130 | Warm-up |  | 5.29 |
| 17 :15 | _0008 | 130 | 130 | isothermal |  |  |

## Annealing to 140 K

| 17:30 | _0009 | 130 | 139.8 | Warm-up |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :45 | _0010 | 139.8 | 140 | Warm-up |  |  |
| 18 :00 | _0011 | 140 | 140 | isothermal |  | 5.27 |

## Annealing to 150 K

| 18:15 | _0012 | 140 |  | Warm-up |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :30 | _0013 |  | 150 | Warm-up |  |  |
| 18 :45 | _0014 | 150 | 150 | isothermal |  | 5.26 |
| 19 :00 | _0015 | 150 | 150 | isothermal |  | 5.25 |
| 19 :15 | _0016 | 150 | 150 | isothermal | ~3.4 * 10-7 | 5.24 |
| 19:30 | _0017 | 150 | 150 | Isothermal | ~3.2 * 10-7 | 5.23 |
| 19:45 | _0018 | 150 | 150 | Isothermal | ~3.1 * 10-7 | 5.23 |
| 20:00 | _0019 | 150 | 150 | Isothermal | ~3.1 * 10-7 | 5.22 |
| 20:15 | _0020 | 150 | 150 | Isothermal | ~3.0 * 10-7 | 5.20 |
| 20:30 | _0021 | 150 | 150 | Isothermal | ~2.9 * 10-7 | 5.19 |
| 20:45 | _0022 | 150 | 150 | Isothermal | ~2.7 * 10-7 | 5.18 |
| 21:00 | _0023 | 150 | 150 | Isothermal | ~2.4 * 10-7 | 5.17 |

## Annealing to 160 K

| 21 : 15 | _0024 |  |  | Warm-up |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 21 :30 | _0025 |  | 160 | Warm-up |  | 5.10 |

# Monday 28th Sep 2020	(VD)

9:18	T = 278.4	P = 5.10 * 10-9

9:18	Cryo turned on

9:25	Laser On

9:44	MCT Detector cooled

10:06	T = 133.2	P = 1.5 10e-9

## Background scan(s) #1

11:02	BG20200928_01

512 scans	res = 1 cm-1	signal = 5.13
p = 7 10-10 mbar	T = 21.7 K

11 :25	Refill gas cell procedure (0.659 Torr initial pressure)

Pump valve close

{...}

Final Pressure : 0.755 Torr

11:30	Head rotated Laser signal = 325.9 mV

## Deposition #6: H2O @ 22K

11:40 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 21.6K

- Initial pressure: 7 * 10-10 mbar

- Initial gas cell pressure: 7.54 Torr

- Deposition pressure: 1.01-1.11 * 10-7 mbar

- Laser signal:  326.2 mV

- Deposition time: 20 min

- pressure after deposition: 9.0 * 10-9 mbar

- final gas cell pressure = 0.636 Torr

11:55	Laser turned off

12:00	Head rotated

12:00	T = 21.7		P = 6.1 * 10-9 mbar

12:00	ASW_2020_09_28_0001

512 scans	1 cm-1 res	signal 5.19
p = 4.5 x 10-9 mbar	T = 21.6 K

12:24	ASW_2020_09_28_0002

512 scans	1 cm-1 res	signal  5.18
p = 1.6 * 10-9	T =  21.6 K

## Annealing to 30 K

| 13 :00 | _0003 | 21.6 | 30 | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 13 :15 | _0004 | 30 | 30 | Warm-up 2 |  |  |
| 13 :30 | _0005 | 30 | 30 | Isothermal 1 |  | 5.14 |
| 13 :45 | _0006 | 30 | 30 | Isothermal 2 |  | 5.14 |

## Annealing to 40 K

| 14 :00 | _0007 |  |  | Warm-up 1 |  | 5.14 |
| --- | --- | --- | --- | --- | --- | --- |
| 14 :15 | _0008 |  |  | Warm-up 2 |  |  |
| 14 :30 | _0009 |  |  | Isothermal 1 |  | 5.13 |
| 14 :45 | _0010 |  |  | Isothermal 2 |  | 5.12 |

## Annealing to 50 K

| 15 :00 | _0011 |  |  | Warm-up 1 |  | 5.12 |
| --- | --- | --- | --- | --- | --- | --- |
| 15 :15 | _0012 |  |  | Warm-up 2 |  | 5.11 |
| 15 :30 | _0013 |  |  | Isothermal 1 |  |  |
| 15 :45 | _0014 |  |  | Isothermal 2 |  | 5.10 |

## Annealing to 60 K

| 16 :00 | _0015 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :15 | _0016 |  |  | Warm-up 2 |  |  |
| 16 :30 | _0017 |  |  | Isothermal 1 |  |  |
| 16 :45 | _0018 | 60 | 60 | Isothermal 2 | 1.3 * 10-9 | 5.07 |

## Annealing to 70 K

| 17 :00 | _0019 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :15 | _0020 |  |  | Warm-up 2 |  |  |
| 17 :30 | _0021 |  |  | Isothermal 1 |  |  |
| 17 :45 | _0022 |  |  | Isothermal 2 |  |  |

## Annealing to 80 K

| 18 :00 | _0023 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :15 | _0024 |  |  | Warm-up 2 |  | 5.06 |
| 18 :30 | _0025 |  |  | Isothermal 1 |  | 5.05 |
| 18 :45 | _0026 |  |  | Isothermal 2 |  |  |

## Annealing to 90 K

| 19 :00 | _0027 |  |  | Warm-up 1 |  | 5.04 |
| --- | --- | --- | --- | --- | --- | --- |
| 19 :15 | _0028 |  |  | Warm-up 2 |  | 5.04 |
| 19 :30 | _0029 |  |  | Isothermal 1 |  | 5.03 |
| 19 :45 | _0030 |  |  | Isothermal 2 | 1.3 * 10-9 | 5.03 |

## Annealing to 100 K

| 20 :00 | _0031 |  |  | Warm-up 1 |  | 5.02 |
| --- | --- | --- | --- | --- | --- | --- |
| 20 :15 | _0032 |  |  | Warm-up 2 |  | 5.02 |
| 20 :30 | _0033 |  |  | Isothermal 1 |  | 5.02 |
| 20 :45 | _0034 |  |  | Isothermal 2 |  |  |

## Annealing to 110 K

| 21 :00 | _0035 |  |  | Warm-up 1 |  | 5.02 |
| --- | --- | --- | --- | --- | --- | --- |
| 21 :15 | _0036 |  |  | Warm-up 2 |  | 5.02 |
| 21 :30 | _0037 |  |  | Isothermal 1 |  | 5.01 |
| 21 :45 | _0038 |  |  | Isothermal 2 |  | 5.01 |

## Annealing to 120 K

| 22 :00 | _0039 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 22 :15 | _0040 |  |  | Warm-up 2 |  |  |
| 22 :30 | _0041 |  |  | Isothermal 1 |  | 5.0 |

22 :45	T = 120K	P = 1.7 * 10-9 mbar

Tuesday 29th of September

10:00 	MCT Detector cooled

10 :26	T = 120K	P = 1.10 * 10-9 mbar

| 10 :30 | _0042 | 120 | 120 | Isothermal 2 | 1.1 * 10 -9 | 4.62 |
| --- | --- | --- | --- | --- | --- | --- |
| 11 :00 | _0043 | 120 | 120 | Isothermal 3 | 1.1 * 10 -9 | 4.67 |
| 11 :15 | _0044 | 120 | 120 | Isothermal 4 | 1.1 * 10 -9 | 4.67 |
| 11 :30 | _0045 | 120 | 120 | Isothermal 4 | 1.1 * 10 -9 | 4.67 |
| 11 :45 | _0046 | 120 | 120 | Isothermal 4 | 1.1 * 10 -9 | 4.67 |

## Annealing to 125 K

| 12 :00 | _0047 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 12 :15 | _0048 |  | 125.0 | Warm-up 2 |  | 4.66 |
| 12 :30 | _0049 |  |  | Isothermal 1 |  |  |
| 12 :45 | _0050 |  |  | Isothermal 2 |  | 4.66 |

## Annealing to 130 K

| 13 :00 | _0051 |  |  | Warm-up 1 |  | 4.66 |
| --- | --- | --- | --- | --- | --- | --- |
| 13 :15 | _0052 |  |  | Warm-up 2 |  |  |
| 13 :30 | _0053 |  |  | Isothermal 1 |  | 4.65 |
| 13 :45 | _0054 |  |  | Isothermal 2 |  | 4.65 |

## Annealing to 135 K

| 14 :00 |  | _0055 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 14 :15 |  | _0056 |  |  | Warm-up 2 |  | 4.65 |
| 14 :30 | _0057 |  |  | Isothermal 1 |  |  |
| 14 :45 | _0058 |  |  | Isothermal 2 |  | 4.65 |

## Annealing to 140 K

| 15 :00 | _0059 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 15 :15 | _0060 |  |  | Warm-up 2 |  | 4.66 |
| 15 :30 | _0061 |  |  | Isothermal 1 |  |  |
| 15 :45 | _0062 |  |  | Isothermal 2 |  | 4.65 |

## Annealing to 145 K

| 16 :00 | _0063 |  |  | Warm-up 1 |  | 4.64 |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :15 | _0064 |  |  | Warm-up 2 |  |  |
| 16 :30 | _0065 |  |  | Isothermal 1 |  |  |
| 16 :45 | _0066 |  |  | Isothermal 2 |  | 4.65 |

## Annealing to 150 K

| 17 :00 | _0067 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :15 | _0068 |  |  | Warm-up 2 |  | 4.65 |
| 17 :30 | _0069 |  |  | Isothermal 1 |  |  |
| 17 :45 | _0070 |  |  | Isothermal 2 | 3.7 * 10-7 | 4.66 |

## Isotherm 150 K

| 18 :00 | _0071 |  |  | Isotherm 3 |  | 4.66 |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :15 | _0072 |  |  | Isotherm 4 | 3.5 * 10-7 | 4.66 |
| 18 :30 | _0073 |  |  | Isotherm 5 | 3.4 * 10-7 | 4.65 |
| 18 :45 | _0074 |  |  | Isotherm 6 | 3.3 * 10-7 | 4.66 |

## Annealing to 155 K

| 19 :00 | _0075 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 19 :15 | _0076 |  |  | Warm-up 2 |  | 4.63 |
| 19 :30 | _0077 |  |  | Warm-up 2 | 7.5 * 10-7 | 4.61 |
| 19 :40 | _0078 |  |  | Warm-up 2 | 6.8 – 4.3 * 10-7 | 4.58 |

## Annealing to 160 K

| 20 :00 | _0079 |  |  | Warm-up 1 |  | 4.59 |
| --- | --- | --- | --- | --- | --- | --- |
| 20 :15 | _0080 |  |  | Warm-up 2 |  | 4.58 |

# Monday 12th Oct 2020	(VD)

9:10	T = 278.4	P = 2.7 * 10-9

9:18	Cryo turned on

11:05	Laser On

11:27	MCT Detector cooled

11:52	T= 21.5	P = 5 * 10-10 mbar

Wrong manipulation while aligning MCT detector … Had to be realigned (box opened …)

No sample produced

# Wednesday 14th Oct 2020	(VD)

10:02	T = 278.6	P = 2.6 * 10-9

10:02	Cryo turned on

10:02	Laser On

10:16	MCT Detector cooled

11:21	T= 21.7 K	P = 5 * 10-10 mbar

## Background scan(s) #1

11:35	BG20201014_01

512 scans	res = 1 cm-1	signal = 3.69
p = 5 10-10 mbar 	T = 21.5 K

11:51	Heater setup to 120K

11:35	BG20201014_02

512 scans	res = 1 cm-1	signal =  3.72
p = 7 10-9 mbar	 T = 119.7 K

## Deposition #7: H2O @ 120K

11:40 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 120 K

- Initial pressure:  * 2.9-9 mbar

- Initial gas cell pressure: 7.51 Torr

- Deposition pressure: 1.01-1.09 * 10-7 mbar

- Laser signal:  78 mV

- Deposition time: 20 min

- pressure after deposition: 8 * 10-9 mbar

- final gas cell pressure = 0.684 Torr

12:56	Laser turned off

12:58	Head rotated

12:58	T = 120K		P = 5.15 * 10-9 mbar

12:58	ASW_2020_09_28_0001

512 scans	1 cm-1 res	signal 3.82
p = 4.9 x 10-9 mbar	T = 120 K

13:27	ASW_2020_09_28_0002

512 scans	1 cm-1 res	signal 3.80
p = 2.2 x 10-9 mbar	T = 120 K

## Annealing to 125 K

| 14 :00 | _0003 | 120 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 14 :15 | _0004 |  |  | Warm-up 2 |  |  |
| 14 :30 | _0005 |  |  | Isothermal 1 |  | 3.79 |

## Annealing to 130 K

| 14 :45 | _0006 | 125 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 15:00 | _0007 |  |  | Warm-up 2 |  | 3.79 |
| 15 :15 | _0008 |  |  | Isothermal 1 |  | 3.78 |

## Annealing to 135 K

| 15 :30 | _0009 | 130 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 15 :45 | _0010 |  |  | Warm-up 2 |  | 3.78 |
| 16 : 00 | _0011 |  |  | Isothermal 1 |  | 3.77 |

## Annealing to 140 K

| 16 :15 | _0012 | 135 |  | Warm-up 1 |  | 3.77 |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :30 | _0013 |  |  | Warm-up 2 |  |  |
| 16 : 45 | _0014 |  |  | Isothermal 1 |  | 3.76 |

## Annealing to 145 K

| 17:00 | _0015 | 140 |  | Warm-up 1 |  | 3.77 |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :15 | _0016 |  |  | Warm-up 2 |  | 3.76 |
| 17 : 30 | _0017 |  |  | Isothermal 1 |  |  |

## Annealing to 150 K

| 17:45 | _0018 | 145 |  | Warm-up 1 |  | 3.76 |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :00 | _0019 |  |  | Warm-up 2 |  |  |
| 18 : 15 | _0020 |  |  | Isothermal 1 |  | 3.73 |
| 18 : 30 | _0021 |  |  | Isothermal 2 |  | 3.74 |

## Annealing to 160 K

| 18 :45 | _0022 | 145 |  | Warm-up 1 |  | 3.72 |
| --- | --- | --- | --- | --- | --- | --- |
| 19 :00 | _0023 |  |  | Warm-up 2 |  | 3.67 |

# Thursday 15th Oct 2020	(VD)

09:30	T = 277	P = 4.2 * 10-9

09:30	Cryo turned on

09:31	Laser On

10:02	T= 186.7 K	P = 2.3 * 10-9 mbar

10:18	MCT Detector cooled

11:21	T= 21.7 K	P = 5 * 10-10 mbar

## Background scan(s) #1

11:35	BG20201015_01

512 scans	res = 1 cm-1	signal = 3.68
p = 8.3 10-10 mbar 	T = 21.2 K

11:46	T setup at 80K

12:11 T = 80K 	P = 1.9 10-9

11:35	BG20201015_02

512 scans	res = 1 cm-1	signal = 3.68
p = 1.8 10-9 mbar 	T = 80.0 K

## Deposition #8: H2O @ 80K

11:40 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 80 K

- Initial pressure:  1.4 * 10-9 mbar

- Initial gas cell pressure: 7.666 Torr

- Deposition pressure:  * 10-7 mbar

- Laser signal:  75 mV

- Deposition time: 20 min

- pressure after deposition: 1.1* 10-8 mbar

- final gas cell pressure = 0.657 Torr

13:27	Laser turned off

13:29	Head rotated

13:29	T = 80.1K		P = 7 * 10-9 mbar

13:29	ASW_2020_10_15_0001

512 scans	1 cm-1 res	signal 3.75
p = 7 x 10-9 mbar	T = 80.1 K

14:19	ASW_2020_10_15_0002

512 scans	1 cm-1 res	signal 3.74
p = 2.1 x 10-9 mbar	T = 80.0 K

## Annealing to 90 K

| 15 :00 | _0003 | 80 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 15 :15 | _0004 |  |  | Warm-up 2 |  | 3.73 |
| 15 :30 | _0005 |  |  | Isothermal 1 |  | 3.73 |

## Annealing to 100 K

| 15 :45 | _0006 | 90 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :00 | _0007 |  |  | Warm-up 2 |  |  |
| 16 :15 | _0008 |  |  | Isothermal 1 |  | 3.72 |

## Annealing to 110 K

| 16 :30 | _0009 | 100 |  | Warm-up 1 |  | 3.72 |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :45 | _0010 |  |  | Warm-up 2 |  |  |
| 17 :00 | _0011 |  |  | Isothermal 1 |  | 3.73 |

## Annealing to 120 K

| 17 :15 | _0012 | 110 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :30 | _0013 |  |  | Warm-up 2 |  | 3.73 |
| 17 :45 | _0014 |  |  | Isothermal 1 |  |  |

## Annealing to 125 K

| 18 :00 | _0015 | 120 |  | Warm-up 1 |  | 3.71 |
| --- | --- | --- | --- | --- | --- | --- |
| 18 :15 | _0016 |  |  | Warm-up 2 |  |  |
| 18 :30 | _0017 |  |  | Isothermal 1 |  | 3.71 |

## Annealing to 130K

| 18 :45 | _0018 | 125 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 19 :00 | _0019 |  |  | Warm-up 2 |  | 3.71 |
| 19 :15 | _0020 |  |  | Isothermal 1 |  | 3.71 |

## Annealing to 135K

| 19 :30 | _0021 | 130 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 19 :45 | _0022 |  |  | Warm-up 2 |  | 3.71 |
| 20 :00 | _0023 |  |  | Isothermal 1 |  | 3.70 |

## Annealing to 140K

| 20 :15 | _0024 | 135 |  | Warm-up 1 |  | 3.70 |
| --- | --- | --- | --- | --- | --- | --- |
| 20 :30 | _0025 |  |  | Warm-up 2 |  | 3.69 |
| 20:45 | _0026 |  |  | Isothermal 1 |  | 3.69 |

## Annealing to 145K

| 21 :00 | _0027 | 140 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 21 :15 | _0028 |  |  | Warm-up 2 |  | 3.70 |
| 21:30 | _0029 |  |  | Isothermal 1 |  | 3.70 |

## Annealing to 150K

| 21 :45 | _0030 | 145 |  | Warm-up 1 |  | 3.69 |
| --- | --- | --- | --- | --- | --- | --- |
| 22 : 00 | _0031 |  |  | Warm-up 2 |  | 3.69 |
| 22:15 | _0032 |  |  | Isothermal 1 | 4.1 * 10-7 | 3.68 |
| 22:30 | _0033 |  |  | Isothermal 2 | 4 * 10-7 | 3.68 |

## Annealing to 160K

| 22 :45 | _0034 | 140 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 23 :00 | _0035 |  |  | Warm-up 2 |  | 3.66 |

# Thursday 22th Oct 2020	(VD)

10:50	T = 278.4	P = 2.2 * 10-9

10:51	Cryo turned on

10:51	Laser On

11:04	T= 241.6 K	P = 1.7 * 10-9 mbar

11:59	MCT Detector cooled

12:00	T= 67.5 K	P = 8 * 10-10 mbar

16;16	T = 21.2 K	P = 5 * 10-10 mbar

# Background scan(s) #1

16:25	BG20201022_01

512 scans	res = 1 cm-1	signal = 3.60
p = 5 10-10 mbar 	T = 21.2 K

## Deposition #9: H2O @ 20K

17:40 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 21.3 K

- Initial pressure:  5 * 10-10 mbar

- Initial gas cell pressure: 7.49 Torr

- Deposition pressure:  * 10-7 mbar

- Laser signal:  59.3 mV

- Deposition time: 20 min

- pressure after deposition: 6.9 * 10-9 mbar

- final gas cell pressure = 6.28 Torr

18:04	Laser turned off

18:06	Head rotated

18:06	T = 21.3K		P = 5.6 * 10-9 mbar

18:08	ASW_2020_10_15_0001

512 scans	1 cm-1 res	signal 3.28
p = 5 x 10-9 mbar	T = 21.3 K

18:30	ASW_2020_10_15_0002

512 scans	1 cm-1 res	signal 3.61
p = 1.6 x 10-9 mbar	T = 21.3 K

## Annealing to 120 K

| 19 :00 | _0003 | 21.2 |  | Warm-up 1 |  | 3.62 |
| --- | --- | --- | --- | --- | --- | --- |
| 19 :15 | _0004 |  |  | Warm-up 2 |  | 3.62 |
| 19 :30 | _0005 |  |  | Isothermal 1 |  | 3.61 |

## Annealing to 130 K

| 19 :45 | _0006 | 120 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 20 :00 | _0007 |  |  | Warm-up 2 |  | 3.61 |
| 20:15 | _0008 |  |  | Isothermal 1 |  | 3.60 |

## Scan 9 – 28 = Isothermal at 130K (MACRO)

Scan 29  09:19 (23-10) : Dummy scan to see  Nothing

09:53	MCT Detector cooled

10:00	ASW_2020_10_15_0030

512 scans	1 cm-1 res	signal 3.63
p = 3.7 x 10-9 mbar	T = 129.9 K	 (Background 01)

## Scan 31 – 34 = Isothermal at 130K (MACRO)

12:45	ASW_2020_10_15_0035

512 scans	1 cm-1 res	signal 3.58
p = 3.6 x 10-9 mbar	T = 130 K	 (Background 01)

## Annealing to 132K

| 13 :00 | _0036 | 130 |  | Warm-up 1 |  | 3.70 |
| --- | --- | --- | --- | --- | --- | --- |
| 13 : 15 | _0037 |  |  | Warm-up 2 |  |  |
| 13:30 | _0038 |  |  | Isothermal 1 |  | 3.69 |

## Annealing to 134K

| 13 :45 | _0039 | 132 |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 14 : 00 | _0040 |  |  | Warm-up 2 |  | 3.69 |
| 14:15 | _0041 |  |  | Isothermal 1 |  |  |

## Annealing to 136K

| 14 :30 | _0042 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 14 : 45 | _0043 |  |  | Warm-up 2 |  | 3.67 |
| 15:00 | _0044 |  |  | Isothermal 1 |  |  |

## Annealing to 138K

| 15 :15 | _0045 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 15 :30 | _0046 |  |  | Warm-up 2 |  | 3.67 |
| 15:45 | _0047 |  |  | Isothermal 1 |  | 3.67 |

## Annealing to 140K

| 16 :00 | _0048 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 16 :15 | _0049 |  |  | Warm-up 2 |  | 3.67 |
| 16:30 | _0050 |  |  | Isothermal 1 |  | 3.67 |

## Annealing to 142K

| 16 :45 | _0051 |  |  | Warm-up 1 |  | 3.66 |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :00 | _0052 |  |  | Warm-up 2 |  | 3.66 |
| 17:15 | _0053 |  |  | Isothermal 1 |  | 3.66 |

## Annealing to 145K

| 17 :30 | _0054 |  |  | Warm-up 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 17 :45 | _0055 |  |  | Warm-up 2 |  | 3.67 |
| 18:00 | _0056 |  |  | Isothermal 1 |  |  |

21:00	ASW_2020_10_15_0057   150K 512 scans	1 cm-1 res

## Scan 58 – 67 = Isothermal at 150K (MACRO)

21:00	ASW_2020_10_15_0068   160K

512 scans	1 cm-1 res	signal 3.64

# Friday 30th Oct 2020	(VD)

11:00	T = 278.4	P = 1.9 * 10-9

11:01	Cryo turned on

11:01	Laser On

11:09	T= 253.7 K	P = 1.7 * 10-9 mbar

11:20	MCT Detector cooled

12:07	T = 74.5		P = 8.2* 10 -10

12:37	T = 21.3		P = 5 * 10-10

12:38	T set up at 130K

13:02	T = 129.8	P = 8.9 10e-9

# Background scan(s) #1

16:25	BG20201030_01

512 scans	res = 1 cm-1	signal = 3.43
p = 6.7 10-10 mbar 	T = 130 K

## Deposition #10: H2O @ 130K

13:47 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 130 K

- Initial pressure:   3.1 * 10-9 mbar

- Initial gas cell pressure: 7.52 Torr

- Deposition pressure:  * 10-7 mbar

- Laser signal:  57.8 mV

- Deposition time: 20 min

- pressure after deposition: 2.7 * 10-8 mbar

- final gas cell pressure = 6.65 Torr

14:09	Laser turned off

14:10	Head rotated

14:10	T = 130K		P = 3.2 * 10-8 mbar

14:25	ASW_2020_10_30_0001

512 scans	1 cm-1 res	signal 3.66
p = 2.9 x 10-8 mbar	T = 130K

Scan 2 – 11 Macro

17 : 20  Scan 12 – 31 Macro

## Saturday 31/10

15 : 13 T = 130	P = 1 * 10-8

15 :37 MCT Detector cooled

16:00	ASW_2020_10_30_0032

512 scans	1 cm-1 res	signal 3.48
p = 9.9 x 10-9 mbar	T = 130K

16 : 15	 Scan 33 – 52 Macro

21 : 10	 Scan 53 – 72 Macro

## Monday 02/1

11 : 05 T = 130	P = 4.9 * 10-9

11 :06 MCT Detector cooled

11:36	ASW_2020_10_30_0073

512 scans	1 cm-1 res	signal 2.84
p = 4.9 x 10-9 mbar	T = 130K

Scan 74 – 93 Macro 130 K (end at 16:47)

## Ramp to 135 K

17:00	ASW_2020_10_30_0094

512 scans	1 cm-1 res	signal 2.94
p = 4.8 x 10-9 mbar	T = 130K

17 :15   Scan 95 – 114 Macro 135 K

03/11  10:49 MCT Detector cooled

11:42	ASW_2020_10_30_0115

512 scans	1 cm-1 res	signal 2.85
p = 1.1 x 10-8 mbar	T = 130K

## Ramp to 140 K

13:00	ASW_2020_10_30_0116

512 scans	1 cm-1 res	signal 2.99
p = 1.1 x 10-8 mbar	T = 135K

13:15	ASW_2020_10_30_0117

512 scans	1 cm-1 res	signal 2.99
p = 3.4 x 10-8 mbar	T = 140K

13:30	ASW_2020_10_30_0118

512 scans	1 cm-1 res	signal 2.99
p = 3.5 x 10-8 mbar	T = 140K

## Ramp to 145 K

13:45	ASW_2020_10_30_0119

512 scans	1 cm-1 res	signal 2.99
p = 3.5 x 10-8 mbar	T = 140K

14:00	ASW_2020_10_30_0120

512 scans	1 cm-1 res	signal 2.98	T = 145K

14:15	ASW_2020_10_30_0121

512 scans	1 cm-1 res	signal 2.96
p = 1.2  x 10-7 mbar	T = 145K

## Ramp to 150 K

14:30	ASW_2020_10_30_0119

512 scans	1 cm-1 res	signal 2.99
p = 3.5 x 10-8 mbar	T = 140K

# 04/11 –

10 :47 MCT detector cooled

11:20	ASW_2020_10_30_0143

512 scans	1 cm-1 res	signal 2.30
p = 6.5 x 10-8 mbar	T = 150K

# Monday 16th Nov 2020	(VD)

10:20	T = 278.4	P = 1.9 * 10-9

10:21	Cryo turned on

11:41	Laser On

11:59	MCT Detector cooled

11:59	T = 21.2		P = 5.2* 10 -10

12:19	T= 21.1		P= 5.2 10-10

# Background scan(s) #1

12:25	BG20201116_01

512 scans	res = 1 cm-1	signal = 3.42
p = 5.3 10-10 mbar 	T = 21.1 K

## Deposition #11: H2O @ 20K

13:47 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 21.1 K

- Initial pressure:   5.7 * 10-10 mbar

- Initial gas cell pressure: 7.72 Torr

- Deposition pressure:  * 10-7 mbar

- Laser signal:  227.6 mV

- Deposition time: 20 min

- pressure after deposition: 1.4 * 10-8 mbar

- final gas cell pressure = 6.46 Torr

14:40	Laser turned off

14:42	Head rotated

14:42	T= 21.1K	P= 8.4 10-9

14:43	ASW_2020_11_16_0001

512 scans	1 cm-1 res	signal 3.04
p = 6.7 x 10-9 mbar	T = 21.1K

15:01	ASW_2020_11_16_0002

512 scans	1 cm-1 res	signal 3.43
p = 2.6 x 10-9 mbar	T = 21.1K

## Annealing to 60K

15:25	ASW_2020_11_16_0003

512 scans	1 cm-1 res	signal 3.43
p = 1.8 x 10-9 mbar	T = 21.1K

15 :40 Scan 4 – 12 = MACRO isotherm at 60 – MACRO FAILED !!!!

16:21	ASW_2020_11_16_0004

512 scans	1 cm-1 res	signal 3.41
p = 2.3 x 10-9 mbar	T = 60.0K

16:35	ASW_2020_11_16_0005

512 scans	1 cm-1 res	signal 3.41
p = 2.1 x 10-9 mbar	T = 60.0K

## Annealing to 70K

17:00	ASW_2020_11_16_0006

512 scans	1 cm-1 res	signal 3.40
p = 1.9 x 10-9 mbar	T = 60k

17:15	ASW_2020_11_16_0007

512 scans	1 cm-1 res	signal 3.40
p = 4.5 x 10-9 mbar	T = 70k

17 :30 Scan 8 – 18 = MACRO isotherm at 70

20:04 Macro finished

## Annealing to 130K

20:10	ASW_2020_11_16_00019

512 scans	1 cm-1 res	signal 3.40
p = 1.3 x 10-9 mbar	T = 70k

20 :25 Scan 20 – 40 = MACRO isotherm at 130

## 17/11

11:25 MCT Detector  cooled

12:05	ASW_2020_11_16_00040

512 scans	1 cm-1 res	signal 3.29
p = 4.8 x 10-9 mbar	T = 130k

12:52	ASW_2020_11_16_00041

512 scans	1 cm-1 res	signal 3.51
p = 4.8 x 10-9 mbar	T = 130k

13:22	ASW_2020_11_16_00042

512 scans	1 cm-1 res	signal 3.51
p = 4.8 x 10-9 mbar	T = 130k

## Annealing to 136K

15:00	ASW_2020_11_16_0043

512 scans	1 cm-1 res	signal 
p = 4.8 x 10-9 mbar	T = 130k

15:15	ASW_2020_11_16_0044

512 scans	1 cm-1 res	signal = 3.46
p = 2.1 x 10-8 mbar	T = 136k

15:30	ASW_2020_11_16_0045

512 scans	1 cm-1 res	signal = 3.46
p = 2.2 x 10-8 mbar	T = 136k

15:45	ASW_2020_11_16_0046

512 scans	1 cm-1 res	signal = 3.46
p = 2.2 x 10-8 mbar	T = 136k

## Annealing to 137K

16:00	ASW_2020_11_16_0047

512 scans	1 cm-1 res	signal 
p = 2.4 x 10-8 mbar	T = 136k

16:15	ASW_2020_11_16_0048

512 scans	1 cm-1 res	signal

16:30	ASW_2020_11_16_0049

512 scans	1 cm-1 res	signal

16:45	ASW_2020_11_16_0050

512 scans	1 cm-1 res	signal

## Annealing to 138K

17:00	ASW_2020_11_16_0051

512 scans	1 cm-1 res	signal

17:15	ASW_2020_11_16_0052

512 scans	1 cm-1 res	signal

17:30	ASW_2020_11_16_0053

512 scans	1 cm-1 res	signal

17:45	ASW_2020_11_16_0054

512 scans	1 cm-1 res	signal

18 : 00 Scan 55 – 64 = MACRO isotherm at 138

20:22 – MACRO Finished T = 133.7 (on Manual) … Don`t know what happened

## Annealing to 140K

20:25	ASW_2020_11_16_0065

512 scans	1 cm-1 res	signal

20:45	ASW_2020_11_16_0066

512 scans	1 cm-1 res	signal

21:05	ASW_2020_11_16_0067

512 scans	1 cm-1 res	signal 3.46

## Annealing to 145K

21:25	ASW_2020_11_16_0068

512 scans	1 cm-1 res	signal 3.44

21:40 Macro 20 scans

## 18/11

11:06	T = 144.8 P = 8.0 * 10-8

11:17	MCT Detector cooled

11:30	ASW_2020_11_16_0089

512 scans	1 cm-1 res	signal 3.25

11:46	Heater turned off / Compressor turned off

# Thursday 19th Nov 2020	(VD)

10:16	T = 278.3	P = 3.1 * 10-9

10:17	Cryo turned on

10:43	Laser On

10:49	MCT Detector cooled

11:59	T = 21.2		P = 5.2* 10 -10

17:03	MCT Detector cooled again

17:34	BG20201119_01

512 scans	res = 1 cm-1	signal = 3.42
p = 9.5 10-10 mbar 	T = 21.1 K

17 :53 Laser turned on (forgot to do it earlier)

## Deposition #11: H2O @ 20K

13:47 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 21.1 K

- Initial pressure:   9.3 * 10-10 mbar

- Initial gas cell pressure: 7.72 Torr

- Deposition pressure:  1* 10-7 mbar

- Laser signal:  227.6 mV

- Deposition time: 21:30 min

- pressure after deposition: 1.4 * 10-8 mbar

- final gas cell pressure = 6.46 Torr

18:33	Laser turned off

18:34	Head rotated

18:34	T= 21.1K	P= 9.5 10-9

18:35	ASW_2020_11_19_0001

512 scans	1 cm-1 res	signal 3.40
p = 8.2 x 10-9 mbar	T = 21.1K

18:xx	ASW_2020_11_19_0002

512 scans	1 cm-1 res	signal 3.38
p = 2.6 x 10-9 mbar	T = 21.1K

## Annealing to 60K

19:10	ASW_2020_11_19_0003

512 scans	1 cm-1 res	signal 3.37
p = 2.4 x 10-9 mbar	T = 21.1K

19 :30 Scan 4 – 10 = MACRO isotherm at 60

## Annealing to 70K

21:05	ASW_2020_11_19_0011

512 scans	1 cm-1 res	signal 3.36

21 :20  MACRO 40 scans isotherm at 70 (up to 51)

## 20/11

13:16	T = 70 P = 1.0 * 10-9

13:17	MCT Detector cooled

14:20	ASW_2020_11_19_0052

512 scans	1 cm-1 res	signal 3.25
p = 1.0 x 10-9 mbar	T = 70K

14:45	ASW_2020_11_19_0053

512 scans	1 cm-1 res	signal 3.42
p = 1.0 x 10-9 mbar	T = 70K

## Annealing to 130K

15:10	ASW_2020_11_16_00054

512 scans	1 cm-1 res	signal 3.42
p = 1.1 x 10-9 mbar

15 :25 Scan 55 – 67 = MACRO isotherm at 130

## Annealing to 135K

18:30  ASW_2020_11_16_00068

512 scans	1 cm-1 res	signal 3.39

18 :45  MACRO 20 isotherm at 135

21/11/20

16:30 MCT Detector cooled

P = 9.8 10e-9	T = 135K

17:10  ASW_2020_11_16_00089

512 scans	1 cm-1 res	signal 3.32

## Annealing to 140K

17:30  ASW_2020_11_16_00090

512 scans	1 cm-1 res	signal 3.50

P= 9.8 e-9 T= 135K

17:45  ASW_2020_11_16_0091

512 scans	1 cm-1 res	signal 3.52

18:00  ASW_2020_11_16_0092

512 scans	1 cm-1 res	signal 3.53

## Annealing to 145K

18:15  ASW_2020_11_16_00093

512 scans	1 cm-1 res	signal 3.54

18 :30  MACRO 100 isotherm at 145 interrupted loop 77

21/11/20

12:21 MCT Detector cooled

P = 5.8 10e-8	T = 145.1 K

13:22  ASW_2020_11_16_0160

512 scans	1 cm-1 res	signal 3.23

T = 145K               P = 4.8 * 10-8

13 :38 Cryo turned off

Heater man 300K and switched off

# Thursday 23th Nov 2020	(VD)

9:57	T = 278.4	P = 3.6 * 10-9

9:58 	T send to 300K (P rise to 10-7 )

10:03	P start to decrease  Heater set up to Man 0

10:05	Cryo turned on

10:05	Laser On

11:13	MCT Detector cooled

11:34 T = 21.2 K		P= 3.6 * 10-9

11: 36	Heater setted up at 40K

12:19	T = 40.0K	P = 3.2 * 10-9

# Background scan(s) #1

12:25	BG20201123_01

512 scans	res = 1 cm-1	signal = 3.59
p = 3.2 10-9 mbar 	T = 40.1 K

## Deposition #12: H2O @ 40K

12:50 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 40.0 K

- Initial pressure:   2.9 * 10-9 mbar

- Initial gas cell pressure: 7.54 Torr

- Deposition pressure:  1* 10-7 mbar

- Laser signal:  246.4 mV

- Deposition time: 20 min

- pressure after deposition: 1.3 * 10-8 mbar

- final gas cell pressure = 6.62 Torr

13:12	Laser turned off

13:15	Head rotated

13:15	ASW_2020_11_23_0001

512 scans	1 cm-1 res	signal 3.82
p = 9.6 x 10-9 mbar	T = 40K

13:46	ASW_2020_11_23_0002

512 scans	1 cm-1 res	signal 3.81
p = 9.6 x 10-9 mbar	T = 40K

## Annealing to 60K

14:05	ASW_2020_11_23_0003

512 scans	1 cm-1 res	signal 
p = 3.0 x 10-9 mbar

14 :20 Scan 4 – 12 = MACRO isotherm at 60

## Annealing to 80K

16:30	ASW_2020_11_23_0013

512 scans	1 cm-1 res	signal 
p = 2.2 x 10-9 mbar

16 :45 Scan 14 – 22 = MACRO isotherm at 80

## Annealing to 100K

19:00	ASW_2020_11_23_0023

512 scans	1 cm-1 res	signal 
p = 2.2 x 10-9 mbar

19 :15 Scan 24 – 32 = MACRO isotherm at 100

## 24/11

09 :43 T = 100K P = 1.4 * 10e-9

10:00	MCT Detector cooled

10:16	ASW_2020_11_23_0033

512 scans	1 cm-1 res	signal 3.65
p = 1.4 x 10-9 mbar T = 100K

10:35	ASW_2020_11_23_0034

512 scans	1 cm-1 res	signal 3.65
p = 1.4 x 10-9 mbar T = 100K

## Annealing to 120K

11:00	ASW_2020_11_23_0035

512 scans	1 cm-1 res	signal 
p = 2.2 x 10-9 mbar

11 :15 Scan 36 – 44 = MACRO isotherm at 120

## Annealing to 130K

11:00	ASW_2020_11_23_0045

512 scans	1 cm-1 res	signal 
p = 2.0 x 10-9 mbar

13 :40 Scan 46 – 54 = MACRO isotherm at 130

## Annealing to 135K

15:50	ASW_2020_11_23_0055

512 scans	1 cm-1 res	signal 
p = 6.5 x 10-9 mbar

16 :05 Scan 56 – 64 = MACRO isotherm at 135

## Annealing to 140K

18:15	ASW_2020_11_23_0065

512 scans	1 cm-1 res	signal 3.83

18 :30 Scan 66 – 74 = MACRO isotherm at 140

20:39 MCT Detector cooled

20:41	ASW_2020_11_23_0075

512 scans	1 cm-1 res	signal 3.69

P = 3.8 * 10-8	T = 140.0

## Annealing to 145K

21:00	ASW_2020_11_23_0076

512 scans	1 cm-1 res	signal 3.37

P= 3.4 10-8	
21 :15  MACRO isotherm at 145     80 scans

## 25/11

13:27	Loop exit at 69 scans   UP TO SCAN 144

P = 2.27 10e-8

13:34 MCT Detector cooled

13:35 Macro 20 scans (isotherm 145)

MACRO 20 scans

18:40 P = 8.5 * 10-9

18:40	ASW_2020_11_23_0165

512 scans	1 cm-1 res	signal 3.78

18 :57 Cryo turned off, heater setted up at 300 (P rise to 3 * 10-7  ice wasn`t fully desorbed)

19:00 P = 4.5 e-7 T = 163.9

19:01 T = 174.4 P = 7.5 e-7

At 177 K P rise to 10-5 and then decrease

# Thursday 26th Nov 2020	(VD)

T = 277.9 P = 4.9 e-9

9:50 Cryo ON

13:56	MCT Detector cooled

14:00 Laser On

14:00 T set up 140K

# Background scan(s) #1

12:25	BG20201126_01

512 scans	res = 1 cm-1	signal = 3.59

## Deposition #12: H2O @ 140K

15:30 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 140.0 K

- Initial pressure:   7.5 * 10-9 mbar

- Initial gas cell pressure: 7.57 Torr

- Deposition pressure:  1* 10-7 mbar

- Laser signal:  237.3 mV

PROBLEM with leak valve … (very thick sample, very quick flow)

15:27	ASW_2020_11_26_0001

512 scans	1 cm-1 res	signal 3.38

15:43	ASW_2020_11_26_0002

512 scans	1 cm-1 res	signal 3.38

16 :00 Heater turned off and Macro 20 scans

20:40 Macro finished

## Annealing to 120K

20:45	ASW_2020_11_26_0023

512 scans	1 cm-1 res	signal 3.34

P = 5 * 10-9	T = 20.7

21:00	ASW_2020_11_26_0024

512 scans	1 cm-1 res	signal 3.34

P = 5.5 * 10-8	T = 119

21:15	ASW_2020_11_26_0025

512 scans	1 cm-1 res	signal 3.34

## Annealing to 140K

21:30	ASW_2020_11_26_0026

512 scans	1 cm-1 res	signal 3.34

21:45	ASW_2020_11_26_0027

512 scans	1 cm-1 res	signal 3.34

22:00	ASW_2020_11_26_0028

512 scans	1 cm-1 res	signal 3.34

22 :18 MCT Detector cooled

22:20	ASW_2020_11_26_0029

512 scans	1 cm-1 res	signal 3.32

## Annealing to 145K

22:37	ASW_2020_11_26_0030

512 scans	1 cm-1 res	signal 3.33

22:52 Macro 60 scans – EXIT AT LOOP 45 up to scan 75

## Annealing to 150K

09:20	ASW_2020_11_26_0076

512 scans	1 cm-1 res	signal 3.33

09:35	ASW_2020_11_26_0077

512 scans	1 cm-1 res	signal 3.32

09:49	ASW_2020_11_26_0078

512 scans	1 cm-1 res	signal 3.32

## Annealing to 160K

10:04	ASW_2020_11_26_0079

512 scans	1 cm-1 res	signal 3.33

10:18	ASW_2020_11_26_0080

512 scans	1 cm-1 res	signal 3.33

10:32	ASW_2020_11_26_0081

512 scans	1 cm-1 res	signal 3.32

10 :48	Cryo turned off

10 :50 Heater 300K

12 :45 P = 3.98 * 10-7

16 :26	P = 1.14 10-7

16 :26 Turbo off

## Friday 27th Nov 2020	(VD)

PAC was switched off last Friday afternoon – for power shut down over the weekend.

T = 277.9

13:35	Rotary on P = 4.96 x 10-3 mbar – switched on and pumped down within minutes

NB – Turbo backing rotary pump has a problem with the on/off switch

13:41	Turbo ON

13:45	1.4x10-6 mbar

13:50 	P = 7x10-7 mbar

13:52	p=5.5x10-7 mbar

## Wednesday 2nd Dec 2020

14:50 T = 278.5	P = 6.4 10-9

Cry-cooling : 14:50 Cry turned on

15:03 T = 247.5 P = 5.8 * 10-9

15:29 T = 170.2 P = 4.7 e-9 mbar

16:00 T = 66.7	P = 3.4 e-9

16:10 T = 21.3 P = 9.3 * 10-10

16:10 	Cryo - off

## Friday 3rd Dec 2020

Ethane gas line installation

09:58	T = 278.0K	P = 5.7 * 10-9 mbar

## Thursday 21st Jan 2021

10:38	T = 278.5	P = 2.3 * 10-8

10:40	Cryo turned on

11:03	T = 210 K		P= 2.05 * 10-8

11:38	T = 103 K		P= 1.55 * 10-8

11:42	Laser On

11:57	T = 121.5 K		P= 2.6 * 10-9

12:08	MCT Detector cooled

12:36 T = T setup to 60K

13:24	T = 60.0K	P = 1.8 * 10-8

# Background scan(s) #1

13:30	BG20210121_01

512 scans	res = 1 cm-1	signal = 3.80
p = 1.8 10-8 mbar 	T = 60 K

Open gas cell valve (amont - bottom) -> P is now 4.9 * 10-9 (don`t know where it was before)

Cf below comparison between background this XP vs 11-16

13:44	BG20210121_02

512 scans	res = 1 cm-1	signal = 3.86
p = 	T = 60 K

## Deposition #14: H2O @ 60K

14:30 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 60.0 K

- Initial pressure:   3.0 * 10-9 mbar

- Initial gas cell pressure: 7.17 Torr

(decrease slightly for some time ~ 10 min)

- Deposition pressure:  1* 10-7 mbar

- Laser signal:  250.1 mV

- Deposition time: 20 min

- pressure after deposition: 1.9 * 10-8 mbar

Open the leak rather than shutt it at the end (few extra material at end of deposition …)

- final gas cell pressure = 5.9 Torr

14:56	Laser turned off

14:58	Head rotated

15:00	ASW_2020_11_23_0001 (bcg 01 used)

512 scans	1 cm-1 res	signal 4.18
p = 9.3 x 10-9 mbar	T = 60K

15:15	ASW_2020_11_23_0002

512 scans	1 cm-1 res	signal 4.20
p = 5.3 x 10-9 mbar	T = 60K

ABORTED

# Monday 25th Nov 2020	(VD)

11:12	T = 278.7	P = 4.0 * 10-9

11:13	Cryo turned on

11:22	T = 255.2	P = 3.7 * 10-9

12:31	T = 21.1	P 1.5 * 10-9

12:56	Laser On

13:05	MCT Detector cooled

13:16	T = 20.8		P = 1.4* 10 -9

13:19 	 T setup to 60K (P rose up to 8 * 10-8)

13:32	T = 59.7	P = 1.7 e-9

# Background scan(s) #1

13:55	BG20210125_01

512 scans	res = 1 cm-1	signal = 4.12
p = 1.7 10-9 mbar 	T = 60 K

## Deposition #14: H2O @ 60K

14:20 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 60.0 K

- Initial pressure:   1.6 * 10-9 mbar

- Initial gas cell pressure: 7.41 Torr

- Deposition pressure:  1 * 10-7 mbar

- Laser signal:  250.0 mV

- Deposition time: 20 min

- pressure after deposition: 1.6 * 10-8 mbar

Open the leak rather than shutt it at the end (few extra material at end of deposition …)

- final gas cell pressure = 6.37 Torr

14:40	Laser turned off

14:42	Head rotated

14:44	ASW_2021_01_25_0001

512 scans	1 cm-1 res	signal 3.96
p = 8.0 x 10-9 mbar	T = 60K

15:04	ASW_2021_01_25_0002

512 scans	1 cm-1 res	signal 3.96
p = 3.4 x 10-9 mbar	T = 60K

## Annealing to 80K

15:30	ASW_2021_01_25_0003

512 scans	1 cm-1 res	signal 3.94
p =

15:45	ASW_2021_01_25_0004

512 scans	1 cm-1 res	signal 3.93
p = 5.5 x 10-9 mbar

16:00	ASW_2021_01_25_0005

512 scans	1 cm-1 res	signal 3.92
p = 4  x 10-9 mbar

16:15	ASW_2021_01_25_0006

512 scans	1 cm-1 res	signal 3.91
p =

## Annealing to 100K

16:30	ASW_2021_01_25_0007

512 scans	1 cm-1 res	signal 3.91
p =

16:45	ASW_2021_01_25_0008

512 scans	1 cm-1 res	signal 3.91
p =

17:00	ASW_2021_01_25_0009

512 scans	1 cm-1 res	signal 3.90
p =

17:15	ASW_2021_01_25_0010

512 scans	1 cm-1 res	signal  3.90
p =

17:30	macro 20 scan

## 26/01

09 :32 T = 100K		 P = 1.7 * 10e-9

09:43	MCT Detector cooled

10:05	ASW_2021_01_25_0031

512 scans	1 cm-1 res	signal  3.79
p = 1.7 * 10-9

10:30	ASW_2021_01_25_0032

512 scans	1 cm-1 res	signal  3.79
p = 1.7 * 10-9

## Annealing to 120K

11:00	ASW_2021_01_25_0033

512 scans	1 cm-1 res	
p = 1.7 * 10-9

11:15	ASW_2021_01_25_0034

512 scans	1 cm-1 res	signal  3.90
p = 5.9 * 10-9

11:30	ASW_2021_01_25_0035

512 scans	1 cm-1 res	signal  
p =

11:45	ASW_2021_01_25_0036

512 scans	1 cm-1 res	signal  3.89
p = 4.2 * 10-9

## Annealing to 130K

12:00	ASW_2021_01_25_0037

512 scans	1 cm-1 res	
p =

12:15	ASW_2021_01_25_0038

512 scans	1 cm-1 res	signal  3.88
p =

12:30	ASW_2021_01_25_0039

512 scans	1 cm-1 res	signal  3.87
p =

12:45	ASW_2021_01_25_0040

512 scans	1 cm-1 res	signal  3.87
p =

## Annealing to 140K

13:00	ASW_2021_01_25_0041

512 scans	1 cm-1 res	
p =

13:15	ASW_2021_01_25_0042

512 scans	1 cm-1 res	signal  3.86
p =

13:30	ASW_2021_01_25_0043

512 scans	1 cm-1 res	signal  3.86
p =

13:45	ASW_2021_01_25_0044

512 scans	1 cm-1 res	
p =

## Annealing to 150K

14:00	ASW_2021_01_25_0045

512 scans	1 cm-1 res	signal  3.86
p = 7.3 * 10-8

14:15	ASW_2021_01_25_0046

512 scans	1 cm-1 res	signal  3.85
p = 4.5 * 10-7

14 :30	Macro 40 scans

## 27/01

11:30 cryo turned off

# Tuesday 09th Feb 2021	(VD)

10:01	Cryo turned on

12:00	Laser On

12:00	MCT Detector cooled

12:57	T = 20.7		P = 5.0 * 10 -10

13:06	BG20210209_01

512 scans	res = 1 cm-1	signal = 3.72
p = 5.0 10-10 mbar 	T = 20.7 K

## Deposition #1: H2O @ 20K

13:47 	2 min @ 1x10-7 mbar H2O

- Initial Temperature: 20.7 K

- Initial pressure:   5.0 * 10-10 mbar

- Initial gas cell pressure: 7.36 Torr

- Deposition pressure:  1* 10-7 mbar

- Laser signal:  325.3 mV

- Deposition time: 2:00 min

- pressure after deposition: 5.8 * 10-9 mbar

- final gas cell pressure = 7.21 Torr

14:08	Head rotated

14:08	T= 20.7K	P= 2.8 10-9

14:10	ASW_2021_02_09_0001

512 scans	1 cm-1 res	signal 3.80
p = 1.8 x 10-9 mbar	T = 20.7K

14:30	ASW_2021_02_09_0002

512 scans	1 cm-1 res	signal 3.80
p = 5.0 x 10-10 mbar	T = 20.6K

14:55	BG20210209_02

512 scans	res = 1 cm-1	signal = 3.72
p = 5.0 10-10 mbar 	T = 20.7 K

## Deposition #2: H2O @ 20K

15:15 	2 min @ 1x10-7 mbar H2O

- Initial Temperature: 20.7 K

- Initial pressure:   5.0 * 10-10 mbar

- Initial gas cell pressure: 7.19 Torr

- Deposition pressure:  1* 10-7 mbar

- Laser signal:  321.9 mV

- Deposition time: 2:00 min

- pressure after deposition: 6.8 * 10-9 mbar

- final gas cell pressure = 7.04 Torr

15:18	Head rotated

15:18	T= 20.7K	P= 3.2 10-9

15:20	ASW_2021_02_09_0003 (BG1)

512 scans	1 cm-1 res	signal 3.80
p = 2.4 x 10-9 mbar	T = 20.7K

15:40	ASW_2021_02_09_0004 (BG2)

512 scans	1 cm-1 res	signal 3.79
p = 5.0 x 10-9 mbar	T = 20.7K

15:58	BG20210209_03

512 scans	res = 1 cm-1	signal = 3.72
p = 5.0 10-10 mbar 	T = 20.6 K

## Deposition #3: H2O @ 20K

16:15 	2 min @ 1x10-7 mbar H2O

- Initial Temperature: 20.7 K

- Initial pressure:   5.0 * 10-10 mbar

- Initial gas cell pressure: 7.04 Torr

- Deposition pressure:  1* 10-7 mbar

- Laser signal:  323.2 mV

- Deposition time: 2:00 min

- pressure after deposition: 6.8 * 10-9 mbar

- final gas cell pressure = 6.89 Torr

16:18	Head rotated

16:20	ASW_2021_02_09_0005 (BG1)

512 scans	1 cm-1 res	signal 3.85
p = 2.5 x 10-9 mbar	T = 20.7K

16:40	ASW_2021_02_09_0006 (BG3)

512 scans	1 cm-1 res	signal 3.85
p = 5.5 x 10-10 mbar	T = 20.7K

16:55	BG20210209_04

512 scans	res = 1 cm-1	signal = 3.72
p = 5.0 10-10 mbar 	T = 20.7 K

## Deposition #4: H2O @ 20K

17:10 	2 min @ 1x10-7 mbar H2O

- Initial Temperature: 20.7 K

- Initial pressure:   5.0 * 10-10 mbar

- Initial gas cell pressure: 6.89 Torr

- Deposition pressure:  1* 10-7 mbar

- Laser signal:  324.9 mV

- Deposition time: 2:00 min

- pressure after deposition: 6.5 * 10-9 mbar

- final gas cell pressure = 6.74 Torr

17:16	Head rotated

17:19	ASW_2021_02_09_0007 (BG1)

512 scans	1 cm-1 res	signal 3.89
p = 2.1 x 10-9 mbar	T = 20.7K

17:35	ASW_2021_02_09_0008 (BG4)

512 scans	1 cm-1 res	signal 3.89
p = 6.3 x 10-9 mbar	T = 20.7K

17:50	BG20210209_05

512 scans	res = 1 cm-1	signal = 3.90
p = 5.0 10-10 mbar 	T = 20.6 K

## Deposition #5: H2O @ 20K

18:10 	2 min @ 1x10-7 mbar H2O

- Initial Temperature: 20.6 K

- Initial pressure:   5.0 * 10-10 mbar

- Initial gas cell pressure: 6.74 Torr

- Deposition pressure:  1* 10-7 mbar

- Laser signal:  324.0 mV

- Deposition time: 2:00 min

- pressure after deposition: 6.5 * 10-9 mbar

- final gas cell pressure = 6.59 Torr

18:13	Head rotated

18:15	ASW_2021_02_09_0009 (BG1)

512 scans	1 cm-1 res	signal 3.91
p = 2.4 x 10-9 mbar	T = 20.7K

18:30	ASW_2021_02_09_0010 (BG5)

512 scans	1 cm-1 res	signal 3.91
p = 7 x 10-10 mbar	T = 20.6K

18 :45 Macro 10 scan (all BG1)

LASER OFF

## 10/02

09:52 MCT Cooled

T = 20.6 P = 5 * 10-10 mbar

10:10	ASW_2021_02_09_0021 (BG1)

512 scans	1 cm-1 res	signal 3.80
p = 5 x 10-10 mbar	T = 20.6K

## Annealing to 40 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 10:30 | _0022 | Warm-up |  |
| 10:45 | _0023 | Iso1 |  |
| 11:00 | _0024 | Iso2 | 3.97 |

## Annealing to 60 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 11:15 | _0025 | Warm-up |  |
| 11:30 | _0026 | Iso1 |  |
| 11:45 | _0027 | Iso2 |  |

## Annealing to 80 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 12:00 | _0028 | Warm-up | 3.93 |
| 12:15 | _0029 | Iso1 | 3.92 |
| 12:30 | _0030 | Iso2 | 3.95 |

## Annealing to 100 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 12:45 | _0031 | Warm-up |  |
| 13:00 | _0032 | Iso1 | 3.93 |
| 13:15 | _0033 | Iso2 | 3.93 |

## Annealing to 120 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 13:30 | _0034 | Warm-up |  |
| 13:45 | _0035 | Iso1 | 3.92 |
| 14:00 | _0036 | Iso2 | 3.92 |

## Annealing to 130 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 14:15 | _0037 | Warm-up |  |
| 14:30 | _0038 | Iso1 | 3.91 |
| 14:45 | _0039 | Iso2 |  |
|  | _0040 | Iso3 |  |
|  | _0041 |  |  |
|  | _0042 | Iso3 |  |

## Annealing to 135 K

_0043

… To input

16 :00 Macro 20 scan 135K

## 11/02

11 :10 MCT cooled T = 135 P 6.9 * 10-9

11:45	ASW_2021_02_09_0064 (BG1)

512 scans	1 cm-1 res	signal 3.90
p = 6.8 x 10-9 mbar	T = 135.0K

13:15	ASW_2021_02_09_0065(BG1)

512 scans	1 cm-1 res	signal 3.50
p = 6.7 x 10-9 mbar	T = 135.0K

16:00	ASW_2021_02_09_0066(BG1)

512 scans	1 cm-1 res	signal 3.50
p = 6.7 x 10-9 mbar	T = 135.0K

## 12/02

12 :00 MCT cooled T = 135 P 5.4 * 10-9

12:30	ASW_2021_02_09_0067(BG1)

512 scans	1 cm-1 res	signal 3.91
p = 5.5 x 10-9 mbar	T = 135.0K

## Annealing to 137 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 12:45 | _0068 | Warm-up | 3.92 |
| 13:00 | _0069 | Iso1 | 3.93 |
| 13:15 | _0070 | Iso2 | 3.92 |

## Annealing to 138 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 13:30 | _0071 | Warm-up |  |
| 13:45 | _0072 | Iso1 | 3.92 |
| 14:00 | _0073 | Iso2 |  |
| 14:15 | _0074 | Iso2 |  |
| 14:30 | _0075 | Iso2 |  |

## Annealing to 140 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 15:00 | _0076 | Warm-up | 3.90 |
| 15:15 | _0077 | Iso1 | 3.89 |
| 15:30 | _0078 | Iso2 | 3.89 |

## Annealing to 145 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 15:45 | _0079 | Warm-up |  |
| 16:00 | _0080 | Iso1 |  |
| 16:15 | _0081 | Iso2 | 3.87 |

## Annealing to 150 K

16:30	ASW_2021_02_09_0082(BG1)

512 scans	1 cm-1 res	signal 3.87
p = 8.5 x 10-8 mbar	T = 145.0K

16 :45 Macro 20 scans

13/02 14 :50 MCT off / Heater off

# Monday 26th April 2021	(VD)

09:53	T = 278.3		P = 6.8 * 10 -10

09:55	Cryo turned on

11:11	T = 21K		P = 5 * 10-10 mbar

11:35	Laser On

11:35	MCT Detector cooled

12:56	T  set up at 140 K

13:32 T = 140K p = 4.5 *10-9 mbar

Problem with spectro, cf labbook (solved)

14:14 T = 140.1K p = 2.2 *10-9 mbar

# Background scan(s) #1

14:20	BG20210426_01

512 scans	res = 1 cm-1	signal = 3.72
p = 2.1 10-9 mbar 	T = 140 K

## Deposition #16: H2O @ 140K

14:20 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 140.0 K

- Initial pressure:   1.7 * 10-9 mbar

- Initial gas cell pressure: 7.28 Torr

- Deposition pressure:  1 * 10-7 mbar

- Laser signal:  325.3 mV

- Deposition time: 20 min

- pressure after deposition: 4.3 * 10-8 mbar

- final gas cell pressure = 7.13 Torr

15:08	Laser turned off

15:08	Head rotated

15:10	ASW_2021_04_26_0001

512 scans	1 cm-1 res	signal 3.70
p = 2.0 x 10-8 mbar	T = 140K

Almost no sample deposited

16:02	BG20210426_02

512 scans	res = 1 cm-1	signal = 3.69
p = 2.1 10-9 mbar 	T = 140 K

16 :10 Laser turned ON

## Deposition #16: H2O @ 140K (1 * 10 -6 deposition pressure)

14:20 	2 min @ 1x10-6 mbar H2O

- Initial Temperature: 140.0 K

- Initial pressure:   1.6 * 10-9 mbar

- Initial gas cell pressure: 7.05 Torr

- Deposition pressure:  1 * 10-6 mbar

- Laser signal:  326.5 mV

- Deposition time: 5 min

- pressure after deposition: 1.3 * 10-7 mbar

- final gas cell pressure = 6.63 Torr

17:57	Laser turned off

17:57	Head rotated

18:00	ASW_2021_04_26_0002

512 scans	1 cm-1 res	signal 3.71
p = 7.5 x 10-8 mbar	T = 140K

18 :15 Macro 20 scans up to scan 22

## 27/04

9:12 MCT Detector cooled

09:30	ASW_2021_04_26_0023

512 scans	1 cm-1 res	signal 3.75
p = 1.9 x 10-8 mbar	T = 140K

09:45	ASW_2021_04_26_0024

512 scans	1 cm-1 res	signal 3.75
p = 1.9 x 10-8 mbar	T = 140K

## Annealing to 150 K

10:00	ASW_2021_04_26_00245

512 scans	1 cm-1 res	signal 3.90
p = 1.9 x 10-8 mbar	T = 140K

10 :15	Macro 5 scans (up to scan 30)

14:11 Heater turned to manual 0 and cryo turned off

# Friday 07th May 2021	(VD)

14:00	Cryo turned on

17:30	Laser On

17:41	MCT Detector cooled

17:54	T = 20.4		P = 5.0 * 10 -10

# Background scan(s) #1

18:20	BG20210507_01

512 scans	res = 1 cm-1	signal = 3.78
p = 5 10-10 mbar 	T = 20.4 K

## Deposition #17: H2O @ 20K

14:20 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 20.4 K

- Initial pressure:   5 * 10-5 mbar

- Initial gas cell pressure: 7.54 Torr

- Deposition pressure:  1 * 10-7 mbar

- Laser signal:  324.2 mV

- Deposition time: 20 min

- pressure after deposition: 1.5 * 10-8 mbar

- final gas cell pressure = 6.35 Torr

19:18	Laser turned off

19:19	Head rotated

19:20	ASW_2021_05_07_0001

512 scans	1 cm-1 res	signal 3.33
p = 5.2 x 10-9 mbar	T = 20.4K

Weired Baseline

19 :35 - Cryo off – START AGAIN TOMORROW (Macro 20 scans)

Clean surface !!

# Friday 08th May 2021	(VD)

12:25 T = 277.9	P = 1.9 * 10-9

12:30 T set up 300K (P rise to 5 * 10-7)

12:56 T = 300 P = 9.2 * 10-9

12:58 heater turn 0 (manual)

13:10 	T = 297.3	P = 6.9 e-9

13:10 Cryo ON

14:05 T = 132.5	P = 2.2 * 10-9

14:16	MCT Detector cooled

# Background scan(s) #1

15:35	BG20210508_01

512 scans	res = 1 cm-1	signal = 3.88
p = 1.6 10-9 mbar 	T = 20.5 K

15 :21	Laser On

## Deposition #17: H2O @ 20K

16:00 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 20.5 K

- Initial pressure:   1.5 * 10-9 mbar

- Initial gas cell pressure: 7.10 Torr

- Deposition pressure:  1 * 10-7 mbar

- Laser signal:  324.3 mV

- Deposition time: 20 min

- pressure after deposition: 1. * 10-8 mbar

- final gas cell pressure = 5.96 Torr

16:21	Laser turned off

16:22	Head rotated

16:23	ASW_2021_05_07_0001

512 scans	1 cm-1 res	signal 3.50
p = 5.4 x 10-9 mbar	T = 20.6K

Still weired background … (will try to turn the chamber more before next scan !)

16:38	ASW_2021_05_07_0002

512 scans	1 cm-1 res	signal 3.52
p = 3.3 x 10-9 mbar	T = 20.5K

16:53	ASW_2021_05_07_0003

512 scans	1 cm-1 res	signal 3.73
p = 2.4 x 10-9 mbar	T = 20.5K

## Annealing to 100 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 17:15 | _0004 | Warm-up | 3.71 |
| 17:30 | _0005 | Iso1 | 3.72 |
| 17:45 | _0006 | Iso2 |  |

Set up heater at 130 before end of scan 0006

17 :58 Annealing to 130K ASW_2021_05_07_0007

Macro 100 scans exit at loop 72 (up to scan 79)

09-05

11:15 MCT Detector cooled

11:15 T = 130.0K	P = 4.6*10e-9

11:43

11:45	ASW_2021_05_07_0080

512 scans	1 cm-1 res	signal 0.75
p = 4.6 x 10-9 mbar	T = 130K

Dodgy Baseline !!! Why ?

(bit of fiddling and ok)

12:15	ASW_2021_05_07_0081

512 scans	1 cm-1 res	signal 3.13
p = 4.6 x 10-9 mbar	T = 130K

12:30	ASW_2021_05_07_0082

512 scans	1 cm-1 res	signal 3.0 (2.73 at sart and then increase …)
p = 4.6 x 10-9 mbar	T = 130K

Macro 3 scans (not performed – baseline dodgy again)

Spectro turned off and on again

# 11/05/2021

16:53	Compressor turned off, heater on manual 0

# Thursday 13th May 2021	(VD)

10:41	MCT Detector cooled

11:28	Cryo turned on

11:28	Laser On

14:06	T = 20.4		P = 8.3 * 10 -10

# Background scan(s) #1

14:10	BG20210513_01

512 scans	res = 1 cm-1	signal = 3.57
p = 8.2 10-10 mbar 	T = 20.4 K

No good background! (blue)

14:24 : cryo stopped

14:25 Heater set up at 300K

15:03 T = 297.5 P start to decrease (from 4.5 * 10-7)

15:05 heater turned off

15:12 T = 290.5 P = 8.0 10-8 Cryo ON

Pressure take long time to go down (probably released lot of crap when warmed up and came down in T too soon)

16:53	T = 20.6 K	P = 7.7 * 10-9

17:00	BG20210513_02

512 scans	res = 1 cm-1	signal = 3.50
p = 7.5 10-9 mbar 	T = 20.5 K

Background even worse …

17 :15 Cryo off / Laser off and start again tomorrow …

14/05/2021

10:34 T = 278.1	P = 3.7 * 10-9

10:35 Cryo ON Laser On

12:10 T = 20.4	P = 2.1 * 10-9

12:15 MCT Detector cooled

13:00	BG20210513_03

512 scans	res = 1 cm-1	signal = 3.77
p = 2.0 10-9 mbar 	T = 20.3 K

## Deposition #18: H2O @ 20K

13:30 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 20.2 K

- Initial pressure:   1.9 * 10-9 mbar

- Initial gas cell pressure: 6.96 Torr

- Deposition pressure:  1 * 10-7 mbar

- Laser signal:  324.5 mV

- Deposition time: 20 min

- pressure after deposition: 2.2 * 10-8 mbar

- final gas cell pressure = 5.86 Torr

14:00	Laser turned off

14:00	Head rotated

14:05	ASW_2021_05_13_0001

512 scans	1 cm-1 res	signal 3.70
p = 8.1 x 10-9 mbar	T = 20.3K

14:40	ASW_2021_05_13_0002

512 scans	1 cm-1 res	signal 3.65
p = 3.1 x 10-9 mbar	T = 20.3K

## Annealing to 70 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 15:00 | _0003 | Warm-up | 3.64 |
| 15:15 | _0004 | Iso1 | 3.63 |
| 15:30 | _0005 | Iso2 | 3.62 |
| 15:45 | _0006 | Iso3 | 3.61 |

## Annealing to 130 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 16:00 | _0007 | Warm-up | 3.62 |
| 16:15 | _0008 | Iso1 | 3.64 |
| 16:30 | _0009 | Iso2 | 3.65 |
| 16:45 | _0010 | Iso3 | 3.66 |
| 17:00 | _0011 | Iso4 | 3.66 |
| 17:15 | _0012 | Iso5 | 3.66 |
| 17:30 | _0013 | Iso6 | 3.66 |
| 17:45 | _0014 | Iso7 | 3.67 |
| 18:00 | _0015 | Iso8 | 3.67 |
| 18:15 | _0016 | Iso9 | 3.66 |
| 18:30 | _0017 | Iso10 | 3.66 |
| 18:45 | _0018 | Iso11 | 3.66 |
| 19:00 | _0019 | Iso12 | 3.66 |
| 19:15 | _0020 | Iso13 | 3.66 |
| 19:30 | _0021 | Iso14 | 3.66 |
| 19:45 | _0022 | Iso15 | 3.66 |
| 20:00 | _0023 | Iso16 | 3.66 |
| 20:15 | _0024 | Iso17 | 3.66 |
| 20:30 | _0025 | Iso18 | 3.66 |
| 20:45 | _0026 | Iso19 | 3.66 |
| 21:00 | _0027 | Iso20 | 3.66 |
| 21:15 | _0028 | Iso21 | 3.66 |

16 :50	P = 1.7 * 10-8

18 :40 P = 1.5 * 10-8

20 :00 P = 9.6 * 10-9

21 :15 P = 7.9 * 10-9

21 :30 MCT detector LN2 refill

21:45	ASW_2021_05_13_0029

512 scans	1 cm-1 res	signal 3.61
p = 7.5 x 10-9 mbar	T = 130K

22 :00 Macro 100 scans

# 15/05/2021

10 :28 Macro in loop 54 (signal = 3.66) T = 130 P = 5.2 * 10-9

11:26 Macro exit loop 58 (signal 3.65) – up to scan 87

11:30 MCT refilled

11:45	ASW_2021_05_13_0088

512 scans	1 cm-1 res	signal 3.65
p = 5.2 x 10-9 mbar	T = 130K

12 :00 Macro 100 scans

15 :30 P = 4.95 * 10-9

18 :36 P = 4.8 * 10-9 (signal 3.61)

19 :59  Loop exited (don`t know number – up to scan 123) – compressed air off !!

20:30	ASW_2021_05_13_0124

512 scans	1 cm-1 res	signal 3.
p = 4.8 x 10-9 mbar	T = 130K

20 :45 Macro 100

16/05

10 :41 Exit at loop 61 (up to scan 208)

Problem with spectro again !!

Computer restarted

10:55	ASW_2021_05_13_0209

512 scans	1 cm-1 res	signal 3.59
p = 4.6 x 10-9 mbar	T = 130K

11:15	ASW_2021_05_13_0210

512 scans	1 cm-1 res	signal 3.68
p = 4.6 x 10-9 mbar	T = 130K

11:30	ASW_2021_05_13_0211

512 scans	1 cm-1 res	signal 3.68
p = 4.6 x 10-9 mbar	T = 130K

11:45	ASW_2021_05_13_0212

512 scans	1 cm-1 res	signal 3.68
p = 4.6 x 10-9 mbar	T = 130K

12:00	ASW_2021_05_13_0213

512 scans	1 cm-1 res	signal 3.67
p = 4.6 x 10-9 mbar	T = 130K

12:15	ASW_2021_05_13_0214

512 scans	1 cm-1 res	signal 3.66
p = 4.6 x 10-9 mbar	T = 130K

12 :30 MCT Detector refilled

12:45	ASW_2021_05_13_0215

512 scans	1 cm-1 res	signal 3.66
p = 4.6 x 10-9 mbar	T = 130K

13 :00 Macro 100 scans

14 :44 compressed air off again …

14 :46 compressed air on again

14 :57 off again (and on)

19:25 Loop exited at 28

19:33 MCT Detector cooled

19:45	ASW_2021_05_13_0244

512 scans	1 cm-1 res	signal 3.61
p = 4.6 x 10-9 mbar	T = 130K

20 :00 Macro 100 scans

17/05/2021

12 :22 Macro exited at loop 71 up to scan 315

12:30 MCT Detector refilled

12:45	ASW_2021_05_13_0316

512 scans	1 cm-1 res	signal 3.65
p = 4.5 x 10-9 mbar	T = 130K

13 :00 Macro 100 scans

16 :04 Compressor off

Exit loop 28 (up to scan )

19:45	ASW_2021_05_13_0345

512 scans	1 cm-1 res	signal 3.65
p = 4.2 x 10-9 mbar	T = 130K

20 :00 Macro 100 scan exit at loop 67 (up to scan )

11:45	ASW_2021_05_13_0413

512 scans	1 cm-1 res	signal 3.69
p = 3.6 x 10-9 mbar	T = 130K

12:15	ASW_2021_05_13_0414

512 scans	1 cm-1 res	signal 3.68
p = 3.5 x 10-9 mbar	T = 130K

Macro 100 scans exit at loop 30 (19 : 15) up to scan 444

19:45	ASW_2021_05_13_0445

512 scans	1 cm-1 res	signal 3.63
p = 3.5 x 10-9 mbar	T = 130K

Macro 100 scans

19/05/2021

11 :21 Exit at loop 67 up to scan

11:45	ASW_2021_05_13_0513

512 scans	1 cm-1 res	signal 3.72
p = 3.0 x 10-9 mbar	T = 130K

12 :00 Macro 100 scans (exit at loop 28 – scan 541)

18:45	ASW_2021_05_13_0542

512 scans	1 cm-1 res	signal 3.63
p = 2.8 x 10-9 mbar	T = 130K

19 :00	Macro 100 scans (exit at loop 71 20/05 11:19)

Up to scan  613 (last one not good signal)

11:45	ASW_2021_05_13_0614

512 scans	1 cm-1 res	signal 3.80
p = 2.5 x 10-9 mbar	T = 130K

12 :00 Macro 100 scans

Exited at loop 28

18:45	ASW_2021_05_13_0643

512 scans	1 cm-1 res	signal 3.79
p = 2.5 x 10-9 mbar	T = 130K

19 :00 Macro 100 scan (exited at loop 74 up to scan 717)

12:45	ASW_2021_05_13_0718

512 scans	1 cm-1 res	signal 3.63
p = 2.4 x 10-9 mbar	T = 130K

13 :00 Macro 100 scans (exit at loop 28 up to scan 746)

19:34 MCT Detector cooled

19:45	ASW_2021_05_13_0747

512 scans	1 cm-1 res	signal 3.73
p = 2.4 x 10-9 mbar	T = 130K

20 :00 Macro 100 scan exit loop 67 up to scan 814

11:37 MCT Detector cooled

11:45	ASW_2021_05_13_0815

512 scans	1 cm-1 res	signal 3.62
p = 2.3 x 10-9 mbar	T = 130K

Macro 100 scans exit at loop 15 up to scan

15:30	ASW_2021_05_13_0831

512 scans	1 cm-1 res	signal 3.63
p = 2.4 x 10-9 mbar	T = 130K

## Warm up to 140K

15:45	ASW_2021_05_13_0832

512 scans	1 cm-1 res	signal 3.63
p = 2.4 x 10-9 mbar	T = 130K

16:00	ASW_2021_05_13_0833

512 scans	1 cm-1 res	signal 3.63
p = 2.5 x 10-8 mbar	T = 140K

16:15	ASW_2021_05_13_0834

512 scans	1 cm-1 res	signal 3.63
p = 2.5 x 10-8 mbar	T = 140K

16:30	ASW_2021_05_13_0835

512 scans	1 cm-1 res	signal 3.63
p = 2.5 x 10-8 mbar	T = 140K

16:45	ASW_2021_05_13_0836

512 scans	1 cm-1 res	signal 3.63
p = 2.4 x 10-8 mbar	T = 140K

17:00	ASW_2021_05_13_0837

512 scans	1 cm-1 res	signal 3.63
p = 2.3 x 10-8 mbar	T = 140K

17:15	ASW_2021_05_13_0838

512 scans	1 cm-1 res	signal 3.62
p = 2.3 x 10-8 mbar	T = 140K

17:30	ASW_2021_05_13_0839

512 scans	1 cm-1 res	signal 3.62
p = 2.3 x 10-8 mbar	T = 140K

17:45	ASW_2021_05_13_0840

512 scans	1 cm-1 res	signal 3.62
p = 2.3 x 10-8 mbar	T = 140K

18:00	ASW_2021_05_13_0841

512 scans	1 cm-1 res	signal 3.62
p = 2.2 x 10-8 mbar	T = 140K

18:15	ASW_2021_05_13_0842

512 scans	1 cm-1 res	signal 3.63
p = 2.2 x 10-8 mbar	T = 140K

## Warm up to 142K

18:30	ASW_2021_05_13_0843

512 scans	1 cm-1 res	signal 3.63
p = 2.4 x 10-8 mbar	T = 140K

18:45	ASW_2021_05_13_0844

512 scans	1 cm-1 res	signal 3.63
p = 3.8 x 10-8 mbar	T = 142 K

19 :00 Macro 8 scans (exit loop 2)

19:30	ASW_2021_05_13_0847

512 scans	1 cm-1 res	signal 3.63
p = 3.7 x 10-8 mbar	T = 142K

19:45	ASW_2021_05_13_0848

512 scans	1 cm-1 res	signal 3.63
p = 3.7 x 10-8 mbar	T = 142K

20:00	ASW_2021_05_13_0849

512 scans	1 cm-1 res	signal 3.63
p = 3.6 x 10-8 mbar	T = 142K

20:15	ASW_2021_05_13_0850

512 scans	1 cm-1 res	signal 3.64
p = 3.6 x 10-8 mbar	T = 142K

20:30	ASW_2021_05_13_0851

512 scans	1 cm-1 res	signal 3.64
p = 3.6 x 10-8 mbar	T = 142K

20:45	ASW_2021_05_13_0852

512 scans	1 cm-1 res	signal 3.64
p = 3.6 x 10-8 mbar	T = 142K

21:00	ASW_2021_05_13_0853

512 scans	1 cm-1 res	signal 3.63
p = 3.6 x 10-8 mbar	T = 142K

## Warm up to 145K

21:15	ASW_2021_05_13_0854

512 scans	1 cm-1 res	signal 3.63
p = 3.6 x 10-8 mbar	T = 142K

21:30	ASW_2021_05_13_0855

512 scans	1 cm-1 res	signal 3.63
p = 7.8 x 10-8 mbar	T = 145K

21:45	ASW_2021_05_13_0856

512 scans	1 cm-1 res	signal 3.63
p = 7.8 x 10-8 mbar	T = 145K

22:00	ASW_2021_05_13_0857

512 scans	1 cm-1 res	signal 3.63
p = 7.8 x 10-8 mbar	T = 145K

22:15	ASW_2021_05_13_0858

512 scans	1 cm-1 res	signal 3.63
p = 7.8 x 10-8 mbar	T = 145K

22:30	ASW_2021_05_13_0859

512 scans	1 cm-1 res	signal 3.62
p = 7.9 x 10-8 mbar	T = 145K

22:45	ASW_2021_05_13_0860

512 scans	1 cm-1 res	signal 3.62
p = 7.9 x 10-8 mbar	T = 145K

23:00	ASW_2021_05_13_0861

512 scans	1 cm-1 res	signal 3.62
p = 7.9 x 10-8 mbar	T = 145K

## Warm up to 150K

23:15	ASW_2021_05_13_0862

512 scans	1 cm-1 res	signal 3.61
p = 7.9 x 10-8 mbar	T = 145K

23:30	ASW_2021_05_13_0863

512 scans	1 cm-1 res	signal 3.61
p = 2.8 x 10-7 mbar	T = 150K

23 :45 Macro 8 scans up to scan 871

23/05/2021

12:20 Cryo off – Heater manual 0

# Monday 24th May 2021	(VD)

10:10	T = 280.5	P = 2.90 * 10-9

10:10	Cryo turned on / Laser On

11:00 T = 134.7	P= 1.8 * 10-9

11:26 MCT Detector cooled

11:26 T = 22.2	P=1.3 * 10-9

12:20 T = 20.3 P= 1.26 * 10-9

# Background scan(s) #1

12:51	BG20210524_01

512 scans	res = 1 cm-1	signal = 3.86
p = 1.2 10-9 mbar 	T = 20.3 K

## Deposition #19: H2O @ 20K

13:15 	20 min @ 1x10-7 mbar H2O

- Initial Temperature: 20.3 K

- Initial pressure:   1.2 * 10-9 mbar

- Initial gas cell pressure: 7.34 Torr

- Deposition pressure:  1 * 10-7 mbar

- Laser signal:  326.3 mV

- Deposition time: 20 min

- pressure after deposition: 1.8 * 10-8 mbar

- final gas cell pressure = 6.22 Torr

13:35	Laser turned off

13:35	Head rotated

13:40	ASW_2021_05_24_0001

512 scans	1 cm-1 res	signal 3.66
p = 6 x 10-9 mbar	T = 20.3K

14:00	ASW_2021_05_24_0002

512 scans	1 cm-1 res	signal 3.66
p = 2.8 x 10-9 mbar	T = 20.3K

14:30	ASW_2021_05_24_0003

512 scans	1 cm-1 res	signal 3.65
p = 2.0 x 10-9 mbar	T = 20.3K

## Annealing to 70 K

| time | File | Type | signal |
| --- | --- | --- | --- |
| 14:45 | _0004 | Warm-up | 3.65 |
| 15:00 | _0005 | Iso1 | 3.63 |
| 15:15 | _0006 | Iso2 | 3.63 |
| 15:30 | _0007 | Iso3 | 3.63 |

## Annealing to 130 K

15:45	ASW_2021_05_24_0008

512 scans	1 cm-1 res	signal 3.66
p = 4.4 x 10-9 mbar	T = 70K

16:00	ASW_2021_05_24_0009

512 scans	1 cm-1 res	signal 3.66
p = 6.5 x 10-8 mbar	T = 133.5K

16:15	ASW_2021_05_24_0010

512 scans	1 cm-1 res	signal 3.67
p = 7.2 x 10-8 mbar	T = 135K

16:30	ASW_2021_05_24_0011

512 scans	1 cm-1 res	signal 3.68
p = 5.5 x 10-8 mbar	T = 135K

16:45	ASW_2021_05_24_0012

512 scans	1 cm-1 res	signal 3.68
p = 3.7 x 10-8 mbar	T = 135K

17:00	ASW_2021_05_24_0013

512 scans	1 cm-1 res	signal 3.68
p = 3.0 x 10-8 mbar	T = 135K

17:15	ASW_2021_05_24_0014

512 scans	1 cm-1 res	signal 3.69
p = 2.6 x 10-8 mbar	T = 135K

17:30	ASW_2021_05_24_0015

512 scans	1 cm-1 res	signal 3.69
p = 2.4 x 10-8 mbar	T = 135K

17:45	ASW_2021_05_24_0016

512 scans	1 cm-1 res	signal 3.69
p = 2.3 x 10-8 mbar	T = 135K

18 :00 MCT Detector refilled

(compressor off …)

18:15	ASW_2021_05_24_0017

512 scans	1 cm-1 res	signal 3.67
p = 2.2 x 10-8 mbar	T = 135K

18 :30 Macro 100 scans (exited at loop 70  up to scan 87)

10:45 MCT Detector cooled

11:00	ASW_2021_05_24_0088

512 scans	1 cm-1 res	signal 3.78
p = 9.6 x 10-9 mbar	T = 135K

11:15	ASW_2021_05_24_0089

512 scans	1 cm-1 res	signal 3.79
p = 9.5 x 10-9 mbar	T = 135K

11 :30 Macro 100 scans exit at loop 29 up to scan 118

18:15 MCT detector refilled

18:30	ASW_2021_05_24_0119

512 scans	1 cm-1 res	signal 3.66
p = 8.7 x 10-9 mbar	T = 135K

18:45	ASW_2021_05_24_0120

512 scans	1 cm-1 res	signal 3.66
p = 8.7 x 10-9 mbar	T = 135K

19 :00 Macro 100 scan exit at loop 68 up to scan 188

11:00	ASW_2021_05_24_0189

512 scans	1 cm-1 res	signal 3.70
p = 8.0 x 10-9 mbar	T = 135K

11:15	ASW_2021_05_24_0190

512 scans	1 cm-1 res	signal 3.69
p = 8.0 x 10-9 mbar	T = 135K

11 :30 Macro 100 exit at loop 30 up to scan 220

18:45	ASW_2021_05_24_0221

512 scans	1 cm-1 res	signal 3.69
p = 8.0 x 10-9 mbar	T = 135K

19 :00 Macro 100 scan exit at loop 71 up to scan 292

27/05

11:45	ASW_2021_05_24_0293

512 scans	1 cm-1 res	signal 3.69
p = 7.5 x 10-9 mbar	T = 135K

12 :00 Macro exited at loop 28 – up to scan 321

18:45	ASW_2021_05_24_0322

512 scans	1 cm-1 res	signal 3.65
p = 7.4 x 10-9 mbar	T = 135K

19 :00 Macro 100 exit at loop 68 up to scan 390

10:50 MCT Detector cooled

11:00	ASW_2021_05_24_0391

512 scans	1 cm-1 res	signal 3.70
p = 7.2 x 10-9 mbar	T = 135K

11:15	ASW_2021_05_24_0392

512 scans	1 cm-1 res	signal 3.70
p = 7.2 x 10-9 mbar	T = 135K

## Warm up to 140K

11:30	ASW_2021_05_13_0393

512 scans	1 cm-1 res	signal 3.70
p = 7.2 x 10-9 mbar	T = 135K

11:45	ASW_2021_05_13_0394

512 scans	1 cm-1 res	signal 3.70
p = 2.5 x 10-8 mbar	T = 140K

12:00	ASW_2021_05_13_0395

512 scans	1 cm-1 res	signal 3.70
p = 2.6 x 10-8 mbar	T = 140K

12:15	ASW_2021_05_13_0396

512 scans	1 cm-1 res	signal 3.70
p = 2.6 x 10-8 mbar	T = 140K

12:30	ASW_2021_05_13_0397

512 scans	1 cm-1 res	signal 3.69
p = 2.6 x 10-8 mbar	T = 140K

12:45	ASW_2021_05_13_0398

512 scans	1 cm-1 res	signal 3.69
p = 2.6 x 10-8 mbar	T = 140K

13:00	ASW_2021_05_13_0399

512 scans	1 cm-1 res	signal 3.69
p = 2.6 x 10-8 mbar	T = 140K

Compressor issues

13:15	ASW_2021_05_13_0400

512 scans	1 cm-1 res	signal 3.73
p = 2.6 x 10-8 mbar	T = 140K

13 :30 Macro 100 exit at loop 17 up to scan 417

17:30	ASW_2021_05_13_0418

512 scans	1 cm-1 res	signal 3.73
p = 2.6 x 10-8 mbar	T = 140K

## Warm up to 145K

17:45	ASW_2021_05_13_0419

512 scans	1 cm-1 res	signal 3.68
p = 2.5 x 10-8 mbar	T = 140K

18:00	ASW_2021_05_13_0420

512 scans	1 cm-1 res	signal 3.67
p = 9 x 10-8 mbar	T = 145K

18:15	ASW_2021_05_13_0421

512 scans	1 cm-1 res	signal 3.67
p = 9.1 x 10-8 mbar	T = 145K

18 :30 MCT Detector refilled

18:45	ASW_2021_05_13_0422

512 scans	1 cm-1 res	signal 3.67
p = 9.3 x 10-8 mbar	T = 145K

19 :00 Macro 100 scans exit at loop 77 up to scan 499

29/05 13: 00 Cryo off, heater off

# End

# General info

## Omnic save

Omnic Autosave file location: C:\My Documents\omnic\Autosave

## Temperature controller

### To set the temperature:

Press [SET] + {RAISE/LOWER] to enter the desire pressure in Kelvin; then press [AUTO]

### To turn off heater (not the controller):

Press [MAN] under ‘Heater’ + [RAISE/LOWER] to reduce the heater voltage to 0.0 V

### PID values

P = 2.5 I = 5.0 D = 0.0

Questions ?

Dangling bonds * 2 or CO2 signature

# DATA Processing

## 16-09

All Isotherms scans (3rd) – smoothed and auto baseline corrected

## 28-09

All Isotherms scans (4th) – smoothed and auto baseline corrected